## README
EA-IRMS correction template
Thermo Flash EA coupled to a Delta/253 IRMS. Takes the Isodat export and returns %C, %N, C:N, delta-13C and delta-15N, with every correction visible as a live formula. ea_irms 1.4.0.
How to use it
1. Open the Raw sheet. Select rows 2 to 1001 and press Delete, then paste your Isodat export starting in cell A2. Do not paste over row 1: the headers there are what every other sheet looks for. The sheet ships loaded with the 2026-09-15 test run as a worked example, so delete that first.
2. Open Setup. Check the standards table against the certificates and against what the sequence actually called things. A material is recognised by the text in column A appearing anywhere in Identifier 1, so "Ace Drift" catches Ace Drift1 through Ace Drift 13. Matching ignores case.
3. Open Work and type sample masses in the yellow column F, in mg. Leave a cell empty to use the mass the export already carries in its Amount column. Without a mass there is no %C, %N or C:N for that row, though the isotope values still come out.
4. Read QC before you read anything else. It flags check standards that missed, samples outside the calibrated size range, blanks that cannot support a blank correction, a reference gas that moved during the run, and positions whose background rose right after an oversized sample.
5. Take the numbers from Results, with their uncertainties. Work is the audit trail behind them and Fits holds every fitted coefficient with its r-squared, residual scatter and, for the corrections, the test that decided whether it was applied.
Colour
Yellow cells with blue text are yours to fill in. Grey columns are machinery and can be ignored. Green columns are the reported values. Everything else is a formula, and overtyping one breaks that row silently, so work on a copy.
What the sheets do
Raw. Paste target, unaltered. Two grey helper columns label each peak record by channel and give it a key.
Setup. Everything you are meant to change: processing switches, the standards table, flag thresholds.
Work. One row per autosampler position. Reads left to right: parsed peaks, then blank, linearity, drift and calibration corrections, then elemental content.
Fits. Every coefficient fitted from this run, each with n, r-squared and residual scatter.
Results. The clean table to copy out or plot from.
QC. Checks and flags. Read this first.
How the corrections work
Blank, then linearity, then drift, then calibration, in that order, each fitted from this run's own standards rather than from a coefficient typed in by hand. Because linearity is fitted after the blank correction, the two cannot double-count a size dependence that the blank already explains.
Calibration regresses accepted values on measured values and applies that line directly. It is never fitted in the other direction and then inverted, which is where the arithmetic usually goes wrong: dividing the intercept by the slope and subtracting it shifts every value by a constant and never applies the slope at all.
Drift is fitted twice, against autosampler position and against the mass-28 background on the reference peaks, and QC reports the residual scatter of both. The default is autosampler position. Residual scatter cannot decide between the two, since both were fitted to the same drift standards, so run labelled check standards and let their offsets decide.
Before trusting any drift correction, read the settling section of QC. What looks like drift is often the instrument still coming to rest: the mass-28 background decays towards a floor over the first several dozen positions and the standards follow it down, then both stop. QC finds the position where the background settles, as the first three readings in a row near the floor, reports the drift slope separately before and after it, and says which pattern the run shows. A spike later in the run is carryover from an oversized sample, not settling, and does not move that position. A straight line in position is the wrong shape for a transient, under-correcting the early positions and over-correcting the late ones. Setup B14 excludes the unsettled positions from the drift fit. The real fix is upstream: wait for the background to reach its floor before the first real sample, and there is no drift left to correct.
QC also lists the scatter of the drift standards after each correction in turn. A correction that does not lower that number is not earning its place in the run you are looking at.
Only corrections the run supports
On Auto, which is the default, a correction is applied only when this run supports it: the blank when at least three blank capsules agree to within 2 per mil, linearity and drift when the fitted slope differs from zero at the significance level in Setup B15 (a two-sided t-test of the slope). That follows the FIRMS Good Practice Guide for IRMS (Dunn and Carter, 2018, section 6.4), which asks that a correction be shown necessary before it is applied. With five or six standards the test has little power, so a slope that fails it is unproven rather than disproven; set the switch to On to apply it regardless. Each drift monitor is also compared against the line through the others, and one that sits far off is flagged rather than removed: if the capsule was the cause, list its position in Setup B16 or B17 and that element's fit leaves it out.
Uncertainty
Results carries a standard uncertainty u for every calibrated value and U = k x u with k from Setup B32. The budget follows section 6.5 of the guide: the 1-sigma certificate uncertainty of each anchor material (Setup columns H and I), the scatter of its replicates in this run, and the repeatability of one sample analysis (Setup B30 and B31, or the pooled anchor scatter when those are empty), each propagated through the calibration line as a Kragten spreadsheet would. It leaves out the blank, linearity and drift corrections and day-to-day reproducibility, so it is a floor. A control chart of an in-house standard is what supplies the rest.
Elemental content comes from a regression of peak area on micrograms of C or N across every standard with both a known mass and a known composition. The intercept of that line is the area blank and is subtracted, so the k-factor is not forced through the origin.
Known limits
Linearity is fitted as a straight line. On the 2026-09-15 run the fitted slope moved 36 percent when the smallest standard was included, so the response is curved across the full size range and a sample far from the standards will carry that curvature as an error. The size flags in QC mark those rows.
A blank correction is only as good as the blanks. Two blank capsules whose delta-15N differ by 22 per mil cannot constrain one, and QC says so rather than letting the correction run quietly. Setting the switch to Auto applies the correction only when at least three blanks agree to within 2 per mil. A sample barely larger than the blank is flagged separately, because there the mass balance amplifies the blank's uncertainty instead of removing it, whatever the blanks look like.
Precision quoted from the drift standards understates the real figure, because those are the material the drift correction was fitted to. QC reports both that number and the scatter of any check standards and repeated unknowns, which is the honest one.
ea_irms 1.4.0. Built at the Wang Stable Isotope Biogeochemistry Lab, Boston College, and free to use and adapt. Cite the version when you publish data corrected with it.

## Raw
Row | Identifier 1 | Identifier 2 | Peak Nr | Amount | Ampl 28 | BGD 28 | Area 28 | Is Ref _ | Ampl 44 | Area 44 | Peak Nr Flash TCD | rArea Flash TCD | d 15N/14N | d 13C/12C | BGD Flash TCD | Channel | Key | grey helper columns, computed
56 | blank capsule |  |  | 0.397 | [6 empty] | 1 | 83626 |  |  | 2123
56 | blank capsule |  | 1 | 0.397 | 673 | 23.3 | 13.036 | 0 |  |  |  |  | -1.246 |  |  | N | 56¦N¦1
56 | blank capsule |  | 2 | 0.397 | 677 | 23.2 | 13.038 | 1 |  |  |  |  | -1.25 |  |  | N | 56¦N¦2
56 | blank capsule |  | 3 | 0.397 | 92 | 25.7 | 2.058 | 0 |  |  |  |  | -6.777 |  |  | N | 56¦N¦3
56 | blank capsule |  | 4 | 0.397 |  |  |  | 1 | 4720 | 89.245 |  |  |  | -25.22 |  | C | 56¦C¦4
56 | blank capsule |  | 5 | 0.397 |  |  |  | 0 | 4685 | 89.236 |  |  |  | -25.445 |  | C | 56¦C¦5
56 | blank capsule |  | 6 | 0.397 |  |  |  | 0 | 4746 | 89.515 |  |  |  | -25.553 |  | C | 56¦C¦6
55 | Ace Drift 13 |  |  | 0.397 | [6 empty] | 1 | 90981 |  |  | 2121
55 | Ace Drift 13 |  |  | 0.397 | [6 empty] | 2 | 683592 |  |  | 2753
55 | Ace Drift 13 |  | 1 | 0.397 | 672 | 23.3 | 13.021 | 0 |  |  |  |  | -1.346 |  |  | N | 55¦N¦1
55 | Ace Drift 13 |  | 2 | 0.397 | 677 | 23.2 | 13.038 | 1 |  |  |  |  | -1.25 |  |  | N | 55¦N¦2
55 | Ace Drift 13 |  | 3 | 0.397 | 1555 | 24 | 28.314 | 0 |  |  |  |  | -6.395 |  |  | N | 55¦N¦3
55 | Ace Drift 13 |  | 4 | 0.397 |  |  |  | 0 | 5635 | 128.087 |  |  |  | -9.436 |  | C | 55¦C¦4
55 | Ace Drift 13 |  | 5 | 0.397 |  |  |  | 1 | 4749 | 89.293 |  |  |  | -25.22 |  | C | 55¦C¦5
55 | Ace Drift 13 |  | 6 | 0.397 |  |  |  | 0 | 4759 | 89.385 |  |  |  | -25.343 |  | C | 55¦C¦6
55 | Ace Drift 13 |  | 7 | 0.397 |  |  |  | 0 | 4753 | 89.531 |  |  |  | -25.405 |  | C | 55¦C¦7
54 | Ace Drift12 |  |  | 0.38 | [6 empty] | 1 | 89197 |  |  | 2122
54 | Ace Drift12 |  |  | 0.38 | [6 empty] | 2 | 655686 |  |  | 2744
54 | Ace Drift12 |  | 1 | 0.38 | 673 | 23 | 13.018 | 0 |  |  |  |  | -1.255 |  |  | N | 54¦N¦1
54 | Ace Drift12 |  | 2 | 0.38 | 677 | 23 | 13.035 | 1 |  |  |  |  | -1.25 |  |  | N | 54¦N¦2
54 | Ace Drift12 |  | 3 | 0.38 | 1494 | 23.8 | 27.037 | 0 |  |  |  |  | -6.297 |  |  | N | 54¦N¦3
54 | Ace Drift12 |  | 4 | 0.38 |  |  |  | 0 | 5383 | 122.82 |  |  |  | -9.406 |  | C | 54¦C¦4
54 | Ace Drift12 |  | 5 | 0.38 |  |  |  | 1 | 4728 | 89.218 |  |  |  | -25.22 |  | C | 54¦C¦5
54 | Ace Drift12 |  | 6 | 0.38 |  |  |  | 0 | 4760 | 89.44 |  |  |  | -25.342 |  | C | 54¦C¦6
54 | Ace Drift12 |  | 7 | 0.38 |  |  |  | 0 | 4761 | 89.366 |  |  |  | -25.42 |  | C | 54¦C¦7
53 | Random | [9 empty] | 1 | 82872 |  |  | 2123
53 | Random |  | 1 |  | 672 | 23.3 | 13.007 | 0 |  |  |  |  | -1.245 |  |  | N | 53¦N¦1
53 | Random |  | 2 |  | 676 | 23.2 | 13.02 | 1 |  |  |  |  | -1.25 |  |  | N | 53¦N¦2
53 | Random |  | 3 |  | 63 | 25.7 | 1.4 | 0 |  |  |  |  | -9.347 |  |  | N | 53¦N¦3
53 | Random |  | 4 |  |  |  |  | 1 | 4723 | 89.184 |  |  |  | -25.22 |  | C | 53¦C¦4
53 | Random |  | 5 |  |  |  |  | 0 | 4706 | 89.357 |  |  |  | -25.479 |  | C | 53¦C¦5
53 | Random |  | 6 |  |  |  |  | 0 | 4735 | 89.215 |  |  |  | -25.543 |  | C | 53¦C¦6
52 | Random | [9 empty] | 1 | 82808 |  |  | 2126
52 | Random | [9 empty] | 2 | 544367 |  |  | 2752
52 | Random |  | 1 |  | 671 | 23.2 | 13.005 | 0 |  |  |  |  | -1.193 |  |  | N | 52¦N¦1
52 | Random |  | 2 |  | 677 | 23.2 | 13.02 | 1 |  |  |  |  | -1.25 |  |  | N | 52¦N¦2
52 | Random |  | 3 |  | 1236 | 24 | 22.447 | 0 |  |  |  |  | -5.883 |  |  | N | 52¦N¦3
52 | Random |  | 4 |  |  |  |  | 0 | 4540 | 101.638 |  |  |  | -9.614 |  | C | 52¦C¦4
52 | Random |  | 5 |  |  |  |  | 1 | 4733 | 89.262 |  |  |  | -25.22 |  | C | 52¦C¦5
52 | Random |  | 6 |  |  |  |  | 0 | 4770 | 89.463 |  |  |  | -25.377 |  | C | 52¦C¦6
52 | Random |  | 7 |  |  |  |  | 0 | 4769 | 89.429 |  |  |  | -25.41 |  | C | 52¦C¦7
51 | Random | [9 empty] | 1 | 74162 |  |  | 2130
51 | Random | [9 empty] | 2 | 383486 |  |  | 2743
51 | Random |  | 1 |  | 671 | 23.3 | 13.001 | 0 |  |  |  |  | -1.183 |  |  | N | 51¦N¦1
51 | Random |  | 2 |  | 676 | 23.3 | 13.019 | 1 |  |  |  |  | -1.25 |  |  | N | 51¦N¦2
51 | Random |  | 3 |  | 855 | 24.1 | 16.082 | 0 |  |  |  |  | -6.109 |  |  | N | 51¦N¦3
51 | Random |  | 4 |  |  |  |  | 0 | 3202 | 71.442 |  |  |  | -9.619 |  | C | 51¦C¦4
51 | Random |  | 5 |  |  |  |  | 1 | 4738 | 89.283 |  |  |  | -25.22 |  | C | 51¦C¦5
51 | Random |  | 6 |  |  |  |  | 0 | 4769 | 89.406 |  |  |  | -25.407 |  | C | 51¦C¦6
51 | Random |  | 7 |  |  |  |  | 0 | 4758 | 89.389 |  |  |  | -25.464 |  | C | 51¦C¦7
50 | Ace Drift11 |  |  | 0.424 | [6 empty] | 1 | 92228 |  |  | 2104
50 | Ace Drift11 |  |  | 0.424 | [6 empty] | 2 | 711683 |  |  | 2757
50 | Ace Drift11 |  | 1 | 0.424 | 671 | 23.4 | 13.006 | 0 |  |  |  |  | -1.258 |  |  | N | 50¦N¦1
50 | Ace Drift11 |  | 2 | 0.424 | 676 | 23.3 | 13.021 | 1 |  |  |  |  | -1.25 |  |  | N | 50¦N¦2
50 | Ace Drift11 |  | 3 | 0.424 | 1623 | 24.2 | 28.692 | 0 |  |  |  |  | -6.161 |  |  | N | 50¦N¦3
50 | Ace Drift11 |  | 4 | 0.424 |  |  |  | 0 | 5892 | 133.29 |  |  |  | -9.488 |  | C | 50¦C¦4
50 | Ace Drift11 |  | 5 | 0.424 |  |  |  | 1 | 4751 | 89.434 |  |  |  | -25.22 |  | C | 50¦C¦5
50 | Ace Drift11 |  | 6 | 0.424 |  |  |  | 0 | 4755 | 89.442 |  |  |  | -25.373 |  | C | 50¦C¦6
50 | Ace Drift11 |  | 7 | 0.424 |  |  |  | 0 | 4756 | 89.506 |  |  |  | -25.413 |  | C | 50¦C¦7
49 | Random | [9 empty] | 1 | 107654 |  |  | 2133
49 | Random | [9 empty] | 2 | 1039471 |  |  | 2769
49 | Random |  | 1 |  | 672 | 23.3 | 13.002 | 0 |  |  |  |  | -1.223 |  |  | N | 49¦N¦1
49 | Random |  | 2 |  | 676 | 23.2 | 13.022 | 1 |  |  |  |  | -1.25 |  |  | N | 49¦N¦2
49 | Random |  | 3 |  | 2404 | 24.1 | 41.648 | 0 |  |  |  |  | -5.971 |  |  | N | 49¦N¦3
49 | Random |  | 4 |  |  |  |  | 0 | 8403 | 195.21 |  |  |  | -9.471 |  | C | 49¦C¦4
49 | Random |  | 5 |  |  |  |  | 1 | 4749 | 89.332 |  |  |  | -25.22 |  | C | 49¦C¦5
49 | Random |  | 6 |  |  |  |  | 0 | 4755 | 89.498 |  |  |  | -25.338 |  | C | 49¦C¦6
49 | Random |  | 7 |  |  |  |  | 0 | 4755 | 89.6 |  |  |  | -25.395 |  | C | 49¦C¦7
48 | Random | [9 empty] | 1 | 89549 |  |  | 2130
48 | Random | [9 empty] | 2 | 680894 |  |  | 2754
48 | Random |  | 1 |  | 672 | 23.1 | 13.007 | 0 |  |  |  |  | -1.244 |  |  | N | 48¦N¦1
48 | Random |  | 2 |  | 676 | 23.1 | 13.022 | 1 |  |  |  |  | -1.25 |  |  | N | 48¦N¦2
48 | Random |  | 3 |  | 1538 | 23.9 | 27.231 | 0 |  |  |  |  | -6.25 |  |  | N | 48¦N¦3
48 | Random |  | 4 |  |  |  |  | 0 | 5640 | 127.336 |  |  |  | -9.587 |  | C | 48¦C¦4
48 | Random |  | 5 |  |  |  |  | 1 | 4725 | 89.404 |  |  |  | -25.22 |  | C | 48¦C¦5
48 | Random |  | 6 |  |  |  |  | 0 | 4759 | 89.363 |  |  |  | -25.357 |  | C | 48¦C¦6
48 | Random |  | 7 |  |  |  |  | 0 | 4763 | 89.585 |  |  |  | -25.4 |  | C | 48¦C¦7
47 | Random | [9 empty] | 1 | 188520 |  |  | 2128
47 | Random |  | 1 |  | 671 | 23.3 | 13.003 | 0 |  |  |  |  | -1.175 |  |  | N | 47¦N¦1
47 | Random |  | 2 |  | 675 | 23.3 | 13.009 | 1 |  |  |  |  | -1.25 |  |  | N | 47¦N¦2
47 | Random |  | 3 |  | 142 | 24.1 | 2.464 | 0 |  |  |  |  | 1.617 |  |  | N | 47¦N¦3
47 | Random |  | 4 |  |  |  |  | 0 | 926 | 20.21 |  |  |  | -10.686 |  | C | 47¦C¦4
47 | Random |  | 5 |  |  |  |  | 1 | 4702 | 89.345 |  |  |  | -25.22 |  | C | 47¦C¦5
47 | Random |  | 6 |  |  |  |  | 0 | 4743 | 89.435 |  |  |  | -25.449 |  | C | 47¦C¦6
47 | Random |  | 7 |  |  |  |  | 0 | 4776 | 89.445 |  |  |  | -25.539 |  | C | 47¦C¦7
46 | Ace Drift10 |  |  | 0.419 | [6 empty] | 1 | 91505 |  |  | 2133
46 | Ace Drift10 |  |  | 0.419 | [6 empty] | 2 | 724362 |  |  | 2757
46 | Ace Drift10 |  | 1 | 0.419 | 671 | 23.1 | 13.003 | 0 |  |  |  |  | -1.263 |  |  | N | 46¦N¦1
46 | Ace Drift10 |  | 2 | 0.419 | 675 | 23.1 | 13.013 | 1 |  |  |  |  | -1.25 |  |  | N | 46¦N¦2
46 | Ace Drift10 |  | 3 | 0.419 | 1652 | 23.9 | 29.047 | 0 |  |  |  |  | -6.231 |  |  | N | 46¦N¦3
46 | Ace Drift10 |  | 4 | 0.419 |  |  |  | 0 | 5971 | 135.61 |  |  |  | -9.407 |  | C | 46¦C¦4
46 | Ace Drift10 |  | 5 | 0.419 |  |  |  | 1 | 4735 | 89.383 |  |  |  | -25.22 |  | C | 46¦C¦5
46 | Ace Drift10 |  | 6 | 0.419 |  |  |  | 0 | 4769 | 89.475 |  |  |  | -25.35 |  | C | 46¦C¦6
46 | Ace Drift10 |  | 7 | 0.419 |  |  |  | 0 | 4750 | 89.43 |  |  |  | -25.403 |  | C | 46¦C¦7
45 | Random | [9 empty] | 1 | 190419 |  |  | 2135
45 | Random |  | 1 |  | 669 | 23.1 | 12.987 | 0 |  |  |  |  | -1.315 |  |  | N | 45¦N¦1
45 | Random |  | 2 |  | 675 | 23.1 | 12.999 | 1 |  |  |  |  | -1.25 |  |  | N | 45¦N¦2
45 | Random |  | 3 |  | 141 | 24 | 2.478 | 0 |  |  |  |  | 1.975 |  |  | N | 45¦N¦3
45 | Random |  | 4 |  |  |  |  | 0 | 945 | 20.462 |  |  |  | -10.909 |  | C | 45¦C¦4
45 | Random |  | 5 |  |  |  |  | 1 | 4711 | 89.291 |  |  |  | -25.22 |  | C | 45¦C¦5
45 | Random |  | 6 |  |  |  |  | 0 | 4720 | 89.297 |  |  |  | -25.444 |  | C | 45¦C¦6
45 | Random |  | 7 |  |  |  |  | 0 | 4753 | 89.331 |  |  |  | -25.521 |  | C | 45¦C¦7
44 | Random | [9 empty] | 1 | 192556 |  |  | 2106
44 | Random |  | 1 |  | 670 | 23.2 | 12.974 | 0 |  |  |  |  | -1.265 |  |  | N | 44¦N¦1
44 | Random |  | 2 |  | 674 | 23.2 | 12.977 | 1 |  |  |  |  | -1.25 |  |  | N | 44¦N¦2
44 | Random |  | 3 |  | 143 | 24 | 2.479 | 0 |  |  |  |  | 1.903 |  |  | N | 44¦N¦3
44 | Random |  | 4 |  |  |  |  | 0 | 954 | 20.814 |  |  |  | -10.725 |  | C | 44¦C¦4
44 | Random |  | 5 |  |  |  |  | 1 | 4710 | 89.289 |  |  |  | -25.22 |  | C | 44¦C¦5
44 | Random |  | 6 |  |  |  |  | 0 | 4730 | 89.244 |  |  |  | -25.415 |  | C | 44¦C¦6
44 | Random |  | 7 |  |  |  |  | 0 | 4764 | 89.347 |  |  |  | -25.479 |  | C | 44¦C¦7
43 | Random | [9 empty] | 1 | 162331 |  |  | 2134
43 | Random |  | 1 |  | 669 | 23.4 | 12.96 | 0 |  |  |  |  | -1.277 |  |  | N | 43¦N¦1
43 | Random |  | 2 |  | 673 | 23.3 | 12.973 | 1 |  |  |  |  | -1.25 |  |  | N | 43¦N¦2
43 | Random |  | 3 |  | 105 | 24.1 | 2.919 | 0 |  |  |  |  | -1.599 |  |  | N | 43¦N¦3
43 | Random |  | 4 |  |  |  |  | 0 | 710 | 15.34 |  |  |  | -10.334 |  | C | 43¦C¦4
43 | Random |  | 5 |  |  |  |  | 1 | 4700 | 89.335 |  |  |  | -25.22 |  | C | 43¦C¦5
43 | Random |  | 6 |  |  |  |  | 0 | 4737 | 89.228 |  |  |  | -25.447 |  | C | 43¦C¦6
43 | Random |  | 7 |  |  |  |  | 0 | 4756 | 89.384 |  |  |  | -25.519 |  | C | 43¦C¦7
42 | Ace Drift9 |  |  | 0.38 | [6 empty] | 1 | 88016 |  |  | 2117
42 | Ace Drift9 |  |  | 0.38 | [6 empty] | 2 | 648873 |  |  | 2766
42 | Ace Drift9 |  | 1 | 0.38 | 668 | 23.2 | 12.946 | 0 |  |  |  |  | -1.09 |  |  | N | 42¦N¦1
42 | Ace Drift9 |  | 2 | 0.38 | 673 | 23.2 | 12.966 | 1 |  |  |  |  | -1.25 |  |  | N | 42¦N¦2
42 | Ace Drift9 |  | 3 | 0.38 | 1469 | 24.1 | 25.931 | 0 |  |  |  |  | -6.24 |  |  | N | 42¦N¦3
42 | Ace Drift9 |  | 4 | 0.38 |  |  |  | 0 | 5413 | 121.336 |  |  |  | -9.51 |  | C | 42¦C¦4
42 | Ace Drift9 |  | 5 | 0.38 |  |  |  | 1 | 4739 | 89.131 |  |  |  | -25.22 |  | C | 42¦C¦5
42 | Ace Drift9 |  | 6 | 0.38 |  |  |  | 0 | 4762 | 89.397 |  |  |  | -25.347 |  | C | 42¦C¦6
42 | Ace Drift9 |  | 7 | 0.38 |  |  |  | 0 | 4751 | 89.418 |  |  |  | -25.432 |  | C | 42¦C¦7
41 | Random | [9 empty] | 1 | 225086 |  |  | 2133
41 | Random |  | 1 |  | 667 | 23.3 | 12.946 | 0 |  |  |  |  | -1.195 |  |  | N | 41¦N¦1
41 | Random |  | 2 |  | 672 | 23.2 | 12.955 | 1 |  |  |  |  | -1.25 |  |  | N | 41¦N¦2
41 | Random |  | 3 |  | 194 | 24.1 | 4.359 | 0 |  |  |  |  | 0.297 |  |  | N | 41¦N¦3
41 | Random |  | 4 |  |  |  |  | 0 | 1238 | 27.265 |  |  |  | -10.813 |  | C | 41¦C¦4
41 | Random |  | 5 |  |  |  |  | 1 | 4718 | 89.222 |  |  |  | -25.22 |  | C | 41¦C¦5
41 | Random |  | 6 |  |  |  |  | 0 | 4755 | 89.394 |  |  |  | -25.415 |  | C | 41¦C¦6
41 | Random |  | 7 |  |  |  |  | 0 | 4768 | 89.408 |  |  |  | -25.475 |  | C | 41¦C¦7
40 | Random | [9 empty] | 1 | 202279 |  |  | 2137
40 | Random |  | 1 |  | 667 | 23.4 | 12.939 | 0 |  |  |  |  | -1.158 |  |  | N | 40¦N¦1
40 | Random |  | 2 |  | 673 | 23.4 | 12.95 | 1 |  |  |  |  | -1.25 |  |  | N | 40¦N¦2
40 | Random |  | 3 |  | 157 | 24.3 | 2.809 | 0 |  |  |  |  | 1.348 |  |  | N | 40¦N¦3
40 | Random |  | 4 |  |  |  |  | 0 | 1041 | 22.965 |  |  |  | -10.776 |  | C | 40¦C¦4
40 | Random |  | 5 |  |  |  |  | 1 | 4695 | 89.181 |  |  |  | -25.22 |  | C | 40¦C¦5
40 | Random |  | 6 |  |  |  |  | 0 | 4732 | 89.325 |  |  |  | -25.44 |  | C | 40¦C¦6
40 | Random |  | 7 |  |  |  |  | 0 | 4763 | 89.298 |  |  |  | -25.525 |  | C | 40¦C¦7
39 | Random | [9 empty] | 1 | 72626 |  |  | 2144
39 | Random | [9 empty] | 2 | 338137 |  |  | 2750
39 | Random |  | 1 |  | 668 | 23.5 | 12.929 | 0 |  |  |  |  | -1.196 |  |  | N | 39¦N¦1
39 | Random |  | 2 |  | 671 | 23.5 | 12.936 | 1 |  |  |  |  | -1.25 |  |  | N | 39¦N¦2
39 | Random |  | 3 |  | 747 | 24.4 | 13.845 | 0 |  |  |  |  | -6.243 |  |  | N | 39¦N¦3
39 | Random |  | 4 |  |  |  |  | 0 | 2836 | 62.657 |  |  |  | -9.567 |  | C | 39¦C¦4
39 | Random |  | 5 |  |  |  |  | 1 | 4705 | 89.175 |  |  |  | -25.22 |  | C | 39¦C¦5
39 | Random |  | 6 |  |  |  |  | 0 | 4751 | 89.375 |  |  |  | -25.386 |  | C | 39¦C¦6
39 | Random |  | 7 |  |  |  |  | 0 | 4755 | 89.36 |  |  |  | -25.468 |  | C | 39¦C¦7
38 | Ace Drift8 |  |  | 0.394 | [6 empty] | 1 | 89275 |  |  | 2145
38 | Ace Drift8 |  |  | 0.394 | [6 empty] | 2 | 684329 |  |  | 2767
38 | Ace Drift8 |  | 1 | 0.394 | 668 | 23.6 | 12.939 | 0 |  |  |  |  | -1.186 |  |  | N | 38¦N¦1
38 | Ace Drift8 |  | 2 | 0.394 | 672 | 23.6 | 12.94 | 1 |  |  |  |  | -1.25 |  |  | N | 38¦N¦2
38 | Ace Drift8 |  | 3 | 0.394 | 1545 | 24.5 | 27.2 | 0 |  |  |  |  | -6.426 |  |  | N | 38¦N¦3
38 | Ace Drift8 |  | 4 | 0.394 |  |  |  | 0 | 5603 | 127.597 |  |  |  | -9.426 |  | C | 38¦C¦4
38 | Ace Drift8 |  | 5 | 0.394 |  |  |  | 1 | 4722 | 89.148 |  |  |  | -25.22 |  | C | 38¦C¦5
38 | Ace Drift8 |  | 6 | 0.394 |  |  |  | 0 | 4750 | 89.294 |  |  |  | -25.377 |  | C | 38¦C¦6
38 | Ace Drift8 |  | 7 | 0.394 |  |  |  | 0 | 4748 | 89.428 |  |  |  | -25.42 |  | C | 38¦C¦7
37 | Random | [9 empty] | 1 | 94726 |  |  | 2145
37 | Random | [9 empty] | 2 | 787133 |  |  | 2768
37 | Random |  | 1 |  | 667 | 23.6 | 12.919 | 0 |  |  |  |  | -1.171 |  |  | N | 37¦N¦1
37 | Random |  | 2 |  | 671 | 23.5 | 12.919 | 1 |  |  |  |  | -1.25 |  |  | N | 37¦N¦2
37 | Random |  | 3 |  | 1790 | 24.5 | 31.222 | 0 |  |  |  |  | -5.902 |  |  | N | 37¦N¦3
37 | Random |  | 4 |  |  |  |  | 0 | 6395 | 146.952 |  |  |  | -9.55 |  | C | 37¦C¦4
37 | Random |  | 5 |  |  |  |  | 1 | 4748 | 89.081 |  |  |  | -25.22 |  | C | 37¦C¦5
37 | Random |  | 6 |  |  |  |  | 0 | 4751 | 89.262 |  |  |  | -25.339 |  | C | 37¦C¦6
37 | Random |  | 7 |  |  |  |  | 0 | 4741 | 89.374 |  |  |  | -25.417 |  | C | 37¦C¦7
36 | Random | [9 empty] | 1 | 84794 |  |  | 2125
36 | Random | [9 empty] | 2 | 575522 |  |  | 2762
36 | Random |  | 1 |  | 666 | 23.6 | 12.89 | 0 |  |  |  |  | -1.155 |  |  | N | 36¦N¦1
36 | Random |  | 2 |  | 670 | 23.6 | 12.903 | 1 |  |  |  |  | -1.25 |  |  | N | 36¦N¦2
36 | Random |  | 3 |  | 1292 | 24.5 | 22.876 | 0 |  |  |  |  | -6.116 |  |  | N | 36¦N¦3
36 | Random |  | 4 |  |  |  |  | 0 | 4715 | 107.199 |  |  |  | -9.529 |  | C | 36¦C¦4
36 | Random |  | 5 |  |  |  |  | 1 | 4727 | 89.177 |  |  |  | -25.22 |  | C | 36¦C¦5
36 | Random |  | 6 |  |  |  |  | 0 | 4744 | 89.3 |  |  |  | -25.39 |  | C | 36¦C¦6
36 | Random |  | 7 |  |  |  |  | 0 | 4744 | 89.245 |  |  |  | -25.437 |  | C | 36¦C¦7
35 | Random | [9 empty] | 1 | 85361 |  |  | 2128
35 | Random | [9 empty] | 2 | 588777 |  |  | 2764
35 | Random |  | 1 |  | 666 | 23.7 | 12.89 | 0 |  |  |  |  | -1.269 |  |  | N | 35¦N¦1
35 | Random |  | 2 |  | 671 | 23.7 | 12.901 | 1 |  |  |  |  | -1.25 |  |  | N | 35¦N¦2
35 | Random |  | 3 |  | 1317 | 24.6 | 23.414 | 0 |  |  |  |  | -6.033 |  |  | N | 35¦N¦3
35 | Random |  | 4 |  |  |  |  | 0 | 4857 | 109.657 |  |  |  | -9.637 |  | C | 35¦C¦4
35 | Random |  | 5 |  |  |  |  | 1 | 4749 | 89.253 |  |  |  | -25.22 |  | C | 35¦C¦5
35 | Random |  | 6 |  |  |  |  | 0 | 4751 | 89.231 |  |  |  | -25.345 |  | C | 35¦C¦6
35 | Random |  | 7 |  |  |  |  | 0 | 4738 | 89.252 |  |  |  | -25.406 |  | C | 35¦C¦7
34 | Ace Drift7 |  |  | 0.399 | [6 empty] | 1 | 89737 |  |  | 2147
34 | Ace Drift7 |  |  | 0.399 | [6 empty] | 2 | 684309 |  |  | 2767
34 | Ace Drift7 |  | 1 | 0.399 | 666 | 24 | 12.89 | 0 |  |  |  |  | -1.205 |  |  | N | 34¦N¦1
34 | Ace Drift7 |  | 2 | 0.399 | 671 | 23.9 | 12.909 | 1 |  |  |  |  | -1.25 |  |  | N | 34¦N¦2
34 | Ace Drift7 |  | 3 | 0.399 | 1544 | 24.8 | 27.2 | 0 |  |  |  |  | -6.213 |  |  | N | 34¦N¦3
34 | Ace Drift7 |  | 4 | 0.399 |  |  |  | 0 | 5661 | 127.619 |  |  |  | -9.469 |  | C | 34¦C¦4
34 | Ace Drift7 |  | 5 | 0.399 |  |  |  | 1 | 4742 | 89.171 |  |  |  | -25.22 |  | C | 34¦C¦5
34 | Ace Drift7 |  | 6 | 0.399 |  |  |  | 0 | 4753 | 89.163 |  |  |  | -25.34 |  | C | 34¦C¦6
34 | Ace Drift7 |  | 7 | 0.399 |  |  |  | 0 | 4752 | 89.215 |  |  |  | -25.394 |  | C | 34¦C¦7
33 | Random | [9 empty] | 1 | 138876 |  |  | 2126
33 | Random | [9 empty] | 2 | 1611823 |  |  | 2799
33 | Random |  | 1 |  | 666 | 23.7 | 12.882 | 0 |  |  |  |  | -1.263 |  |  | N | 33¦N¦1
33 | Random |  | 2 |  | 670 | 23.7 | 12.898 | 1 |  |  |  |  | -1.25 |  |  | N | 33¦N¦2
33 | Random |  | 3 |  | 3810 | 24.7 | 64.233 | 0 |  |  |  |  | -5.821 |  |  | N | 33¦N¦3
33 | Random |  | 4 |  |  |  |  | 0 | 12392 | 302.204 |  |  |  | -9.333 |  | C | 33¦C¦4
33 | Random |  | 5 |  |  |  |  | 1 | 4749 | 89.024 |  |  |  | -25.22 |  | C | 33¦C¦5
33 | Random |  | 6 |  |  |  |  | 0 | 4737 | 89.084 |  |  |  | -25.314 |  | C | 33¦C¦6
33 | Random |  | 7 |  |  |  |  | 0 | 4722 | 89.182 |  |  |  | -25.356 |  | C | 33¦C¦7
32 | Random | [9 empty] | 1 | 73161 |  |  | 2128
32 | Random | [9 empty] | 2 | 304515 |  |  | 2726
32 | Random |  | 1 |  | 664 | 23.9 | 12.876 | 0 |  |  |  |  | -1.253 |  |  | N | 32¦N¦1
32 | Random |  | 2 |  | 669 | 23.9 | 12.888 | 1 |  |  |  |  | -1.25 |  |  | N | 32¦N¦2
32 | Random |  | 3 |  | 654 | 24.9 | 12.229 | 0 |  |  |  |  | -6.39 |  |  | N | 32¦N¦3
32 | Random |  | 4 |  |  |  |  | 0 | 2542 | 56.107 |  |  |  | -9.702 |  | C | 32¦C¦4
32 | Random |  | 5 |  |  |  |  | 1 | 4722 | 89.087 |  |  |  | -25.22 |  | C | 32¦C¦5
32 | Random |  | 6 |  |  |  |  | 0 | 4743 | 89.221 |  |  |  | -25.392 |  | C | 32¦C¦6
32 | Random |  | 7 |  |  |  |  | 0 | 4732 | 89.135 |  |  |  | -25.451 |  | C | 32¦C¦7
31 | Random | [9 empty] | 1 | 87030 |  |  | 2153
31 | Random | [9 empty] | 2 | 656738 |  |  | 2757
31 | Random |  | 1 |  | 665 | 24 | 12.862 | 0 |  |  |  |  | -1.151 |  |  | N | 31¦N¦1
31 | Random |  | 2 |  | 668 | 23.9 | 12.874 | 1 |  |  |  |  | -1.25 |  |  | N | 31¦N¦2
31 | Random |  | 3 |  | 1470 | 25 | 25.209 | 0 |  |  |  |  | -5.844 |  |  | N | 31¦N¦3
31 | Random |  | 4 |  |  |  |  | 0 | 5416 | 121.954 |  |  |  | -9.642 |  | C | 31¦C¦4
31 | Random |  | 5 |  |  |  |  | 1 | 4714 | 89.145 |  |  |  | -25.22 |  | C | 31¦C¦5
31 | Random |  | 6 |  |  |  |  | 0 | 4742 | 89.011 |  |  |  | -25.396 |  | C | 31¦C¦6
31 | Random |  | 7 |  |  |  |  | 0 | 4739 | 89.165 |  |  |  | -25.411 |  | C | 31¦C¦7
30 | Ace Drift6 |  |  | 0.418 | [6 empty] | 1 | 90915 |  |  | 2152
30 | Ace Drift6 |  |  | 0.418 | [6 empty] | 2 | 724909 |  |  | 2760
30 | Ace Drift6 |  | 1 | 0.418 | 664 | 24.2 | 12.852 | 0 |  |  |  |  | -1.167 |  |  | N | 30¦N¦1
30 | Ace Drift6 |  | 2 | 0.418 | 668 | 24.1 | 12.865 | 1 |  |  |  |  | -1.25 |  |  | N | 30¦N¦2
30 | Ace Drift6 |  | 3 | 0.418 | 1626 | 25.1 | 27.844 | 0 |  |  |  |  | -6.19 |  |  | N | 30¦N¦3
30 | Ace Drift6 |  | 4 | 0.418 |  |  |  | 0 | 5934 | 134.66 |  |  |  | -9.478 |  | C | 30¦C¦4
30 | Ace Drift6 |  | 5 | 0.418 |  |  |  | 1 | 4734 | 88.959 |  |  |  | -25.22 |  | C | 30¦C¦5
30 | Ace Drift6 |  | 6 | 0.418 |  |  |  | 0 | 4746 | 89.154 |  |  |  | -25.352 |  | C | 30¦C¦6
30 | Ace Drift6 |  | 7 | 0.418 |  |  |  | 0 | 4736 | 89.138 |  |  |  | -25.409 |  | C | 30¦C¦7
29 | Random | [9 empty] | 1 | 121041 |  |  | 2133
29 | Random | [9 empty] | 2 | 1287398 |  |  | 2794
29 | Random |  | 1 |  | 664 | 24.1 | 12.851 | 0 |  |  |  |  | -1.3 |  |  | N | 29¦N¦1
29 | Random |  | 2 |  | 668 | 24.1 | 12.865 | 1 |  |  |  |  | -1.25 |  |  | N | 29¦N¦2
29 | Random |  | 3 |  | 3000 | 25.1 | 50.397 | 0 |  |  |  |  | -5.694 |  |  | N | 29¦N¦3
29 | Random |  | 4 |  |  |  |  | 0 | 10125 | 240.674 |  |  |  | -9.519 |  | C | 29¦C¦4
29 | Random |  | 5 |  |  |  |  | 1 | 4735 | 89.15 |  |  |  | -25.22 |  | C | 29¦C¦5
29 | Random |  | 6 |  |  |  |  | 0 | 4745 | 89.035 |  |  |  | -25.35 |  | C | 29¦C¦6
29 | Random |  | 7 |  |  |  |  | 0 | 4731 | 89.128 |  |  |  | -25.419 |  | C | 29¦C¦7
28 | Random | [9 empty] | 1 | 84886 |  |  | 2159
28 | Random | [9 empty] | 2 | 386669 |  |  | 2757
28 | Random |  | 1 |  | 665 | 24.2 | 12.864 | 0 |  |  |  |  | -1.275 |  |  | N | 28¦N¦1
28 | Random |  | 2 |  | 668 | 24.1 | 12.871 | 1 |  |  |  |  | -1.25 |  |  | N | 28¦N¦2
28 | Random |  | 3 |  | 1325 | 25.3 | 22.944 | 0 |  |  |  |  | -6.26 |  |  | N | 28¦N¦3
28 | Random |  | 4 |  |  |  |  | 0 | 3218 | 71.395 |  |  |  | -9.387 |  | C | 28¦C¦4
28 | Random |  | 5 |  |  |  |  | 1 | 4704 | 89.121 |  |  |  | -25.22 |  | C | 28¦C¦5
28 | Random |  | 6 |  |  |  |  | 0 | 4744 | 89.135 |  |  |  | -25.399 |  | C | 28¦C¦6
28 | Random |  | 7 |  |  |  |  | 0 | 4743 | 89.096 |  |  |  | -25.47 |  | C | 28¦C¦7
27 | Random | [9 empty] | 1 | 119978 |  |  | 2135
27 | Random | [9 empty] | 2 | 311507 |  |  | 2753
27 | Random |  | 1 |  | 662 | 24.3 | 12.84 | 0 |  |  |  |  | -1.175 |  |  | N | 27¦N¦1
27 | Random |  | 2 |  | 667 | 24.2 | 12.85 | 1 |  |  |  |  | -1.25 |  |  | N | 27¦N¦2
27 | Random |  | 3 |  | 2747 | 25.4 | 47.171 | 0 |  |  |  |  | 0.793 |  |  | N | 27¦N¦3
27 | Random |  | 4 |  |  |  |  | 0 | 2617 | 57.484 |  |  |  | -23.581 |  | C | 27¦C¦4
27 | Random |  | 5 |  |  |  |  | 1 | 4711 | 89.023 |  |  |  | -25.22 |  | C | 27¦C¦5
27 | Random |  | 6 |  |  |  |  | 0 | 4752 | 89.128 |  |  |  | -25.375 |  | C | 27¦C¦6
27 | Random |  | 7 |  |  |  |  | 0 | 4738 | 89.095 |  |  |  | -25.416 |  | C | 27¦C¦7
26 | Ace Drift5 |  |  | 0.4 | [6 empty] | 1 | 90373 |  |  | 2156
26 | Ace Drift5 |  |  | 0.4 | [6 empty] | 2 | 703004 |  |  | 2766
26 | Ace Drift5 |  | 1 | 0.4 | 662 | 24.4 | 12.823 | 0 |  |  |  |  | -1.344 |  |  | N | 26¦N¦1
26 | Ace Drift5 |  | 2 | 0.4 | 667 | 24.3 | 12.833 | 1 |  |  |  |  | -1.25 |  |  | N | 26¦N¦2
26 | Ace Drift5 |  | 3 | 0.4 | 1588 | 25.5 | 27.189 | 0 |  |  |  |  | -6.094 |  |  | N | 26¦N¦3
26 | Ace Drift5 |  | 4 | 0.4 |  |  |  | 0 | 5756 | 130.395 |  |  |  | -9.431 |  | C | 26¦C¦4
26 | Ace Drift5 |  | 5 | 0.4 |  |  |  | 1 | 4727 | 88.96 |  |  |  | -25.22 |  | C | 26¦C¦5
26 | Ace Drift5 |  | 6 | 0.4 |  |  |  | 0 | 4742 | 89.096 |  |  |  | -25.317 |  | C | 26¦C¦6
26 | Ace Drift5 |  | 7 | 0.4 |  |  |  | 0 | 4738 | 89.137 |  |  |  | -25.373 |  | C | 26¦C¦7
25 | Random | [9 empty] | 1 | 121446 |  |  | 2141
25 | Random | [9 empty] | 2 | 318478 |  |  | 2740
25 | Random |  | 1 |  | 661 | 24.4 | 12.802 | 0 |  |  |  |  | -1.184 |  |  | N | 25¦N¦1
25 | Random |  | 2 |  | 666 | 24.4 | 12.822 | 1 |  |  |  |  | -1.25 |  |  | N | 25¦N¦2
25 | Random |  | 3 |  | 2784 | 25.6 | 47.866 | 0 |  |  |  |  | 39.073 |  |  | N | 25¦N¦3
25 | Random |  | 4 |  |  |  |  | 0 | 2628 | 58.45 |  |  |  | 15.888 |  | C | 25¦C¦4
25 | Random |  | 5 |  |  |  |  | 1 | 4711 | 88.938 |  |  |  | -25.22 |  | C | 25¦C¦5
25 | Random |  | 6 |  |  |  |  | 0 | 4725 | 89.013 |  |  |  | -25.384 |  | C | 25¦C¦6
25 | Random |  | 7 |  |  |  |  | 0 | 4743 | 89.147 |  |  |  | -25.469 |  | C | 25¦C¦7
24 | Random | [9 empty] | 1 | 88022 |  |  | 2157
24 | Random | [9 empty] | 2 | 662446 |  |  | 2765
24 | Random |  | 1 |  | 661 | 24.5 | 12.774 | 0 |  |  |  |  | -1.356 |  |  | N | 24¦N¦1
24 | Random |  | 2 |  | 666 | 24.4 | 12.8 | 1 |  |  |  |  | -1.25 |  |  | N | 24¦N¦2
24 | Random |  | 3 |  | 1482 | 25.7 | 25.453 | 0 |  |  |  |  | -6.042 |  |  | N | 24¦N¦3
24 | Random |  | 4 |  |  |  |  | 0 | 5413 | 122.649 |  |  |  | -9.519 |  | C | 24¦C¦4
24 | Random |  | 5 |  |  |  |  | 1 | 4717 | 88.866 |  |  |  | -25.22 |  | C | 24¦C¦5
24 | Random |  | 6 |  |  |  |  | 0 | 4734 | 89.036 |  |  |  | -25.366 |  | C | 24¦C¦6
24 | Random |  | 7 |  |  |  |  | 0 | 4750 | 89.188 |  |  |  | -25.416 |  | C | 24¦C¦7
23 | Random | [9 empty] | 1 | 136064 |  |  | 2167
23 | Random | [9 empty] | 2 | 99729 |  |  | 2748
23 | Random |  | 1 |  | 660 | 24.7 | 12.772 | 0 |  |  |  |  | -1.259 |  |  | N | 23¦N¦1
23 | Random |  | 2 |  | 664 | 24.7 | 12.792 | 1 |  |  |  |  | -1.25 |  |  | N | 23¦N¦2
23 | Random |  | 3 |  | 3433 | 26 | 59.328 | 0 |  |  |  |  | -3.251 |  |  | N | 23¦N¦3
23 | Random |  | 4 |  |  |  |  | 0 | 825 | 18.131 |  |  |  | -18.113 |  | C | 23¦C¦4
23 | Random |  | 5 |  |  |  |  | 1 | 4687 | 88.858 |  |  |  | -25.22 |  | C | 23¦C¦5
23 | Random |  | 6 |  |  |  |  | 0 | 4729 | 88.916 |  |  |  | -25.458 |  | C | 23¦C¦6
23 | Random |  | 7 |  |  |  |  | 0 | 4750 | 89.065 |  |  |  | -25.507 |  | C | 23¦C¦7
22 | Ace Drift4 |  |  | 0.402 | [6 empty] | 1 | 90232 |  |  | 2159
22 | Ace Drift4 |  |  | 0.402 | [6 empty] | 2 | 704695 |  |  | 2775
22 | Ace Drift4 |  | 1 | 0.402 | 661 | 24.8 | 12.794 | 0 |  |  |  |  | -1.244 |  |  | N | 22¦N¦1
22 | Ace Drift4 |  | 2 | 0.402 | 665 | 24.8 | 12.8 | 1 |  |  |  |  | -1.25 |  |  | N | 22¦N¦2
22 | Ace Drift4 |  | 3 | 0.402 | 1591 | 26.1 | 27.288 | 0 |  |  |  |  | -5.874 |  |  | N | 22¦N¦3
22 | Ace Drift4 |  | 4 | 0.402 |  |  |  | 0 | 5751 | 130.484 |  |  |  | -9.538 |  | C | 22¦C¦4
22 | Ace Drift4 |  | 5 | 0.402 |  |  |  | 1 | 4724 | 88.908 |  |  |  | -25.22 |  | C | 22¦C¦5
22 | Ace Drift4 |  | 6 | 0.402 |  |  |  | 0 | 4747 | 89.075 |  |  |  | -25.363 |  | C | 22¦C¦6
22 | Ace Drift4 |  | 7 | 0.402 |  |  |  | 0 | 4732 | 88.99 |  |  |  | -25.416 |  | C | 22¦C¦7
21 | Random | [9 empty] | 1 | 69973 |  |  | 2163
21 | Random | [9 empty] | 2 | 268309 |  |  | 2736
21 | Random |  | 1 |  | 661 | 25.1 | 12.79 | 0 |  |  |  |  | -1.121 |  |  | N | 21¦N¦1
21 | Random |  | 2 |  | 664 | 25.1 | 12.791 | 1 |  |  |  |  | -1.25 |  |  | N | 21¦N¦2
21 | Random |  | 3 |  | 577 | 26.4 | 10.11 | 0 |  |  |  |  | -5.969 |  |  | N | 21¦N¦3
21 | Random |  | 4 |  |  |  |  | 0 | 2216 | 49.134 |  |  |  | -9.66 |  | C | 21¦C¦4
21 | Random |  | 5 |  |  |  |  | 1 | 4690 | 88.868 |  |  |  | -25.22 |  | C | 21¦C¦5
21 | Random |  | 6 |  |  |  |  | 0 | 4730 | 88.898 |  |  |  | -25.441 |  | C | 21¦C¦6
21 | Random |  | 7 |  |  |  |  | 0 | 4732 | 89.004 |  |  |  | -25.51 |  | C | 21¦C¦7
20 | Random | [9 empty] | 1 | 92840 |  |  | 2164
20 | Random | [9 empty] | 2 | 756463 |  |  | 2772
20 | Random |  | 1 |  | 659 | 25.4 | 12.765 | 0 |  |  |  |  | -1.305 |  |  | N | 20¦N¦1
20 | Random |  | 2 |  | 664 | 25.3 | 12.782 | 1 |  |  |  |  | -1.25 |  |  | N | 20¦N¦2
20 | Random |  | 3 |  | 1705 | 26.7 | 29.179 | 0 |  |  |  |  | -5.819 |  |  | N | 20¦N¦3
20 | Random |  | 4 |  |  |  |  | 0 | 6159 | 139.977 |  |  |  | -9.593 |  | C | 20¦C¦4
20 | Random |  | 5 |  |  |  |  | 1 | 4733 | 88.8 |  |  |  | -25.22 |  | C | 20¦C¦5
20 | Random |  | 6 |  |  |  |  | 0 | 4730 | 88.838 |  |  |  | -25.373 |  | C | 20¦C¦6
20 | Random |  | 7 |  |  |  |  | 0 | 4719 | 88.935 |  |  |  | -25.417 |  | C | 20¦C¦7
19 | Ace Bypass |  |  | 0.454 | [6 empty] | 1 | 94190 |  |  | 2166
19 | Ace Bypass |  |  | 0.454 | [6 empty] | 2 | 775660 |  |  | 2775
19 | Ace Bypass |  | 1 | 0.454 | 658 | 25.5 | 12.744 | 0 |  |  |  |  | -1.131 |  |  | N | 19¦N¦1
19 | Ace Bypass |  | 2 | 0.454 | 663 | 25.4 | 12.752 | 1 |  |  |  |  | -1.25 |  |  | N | 19¦N¦2
19 | Ace Bypass |  | 3 | 0.454 | 1745 | 26.9 | 29.846 | 0 |  |  |  |  | -5.941 |  |  | N | 19¦N¦3
19 | Ace Bypass |  | 4 | 0.454 |  |  |  | 0 | 6256 | 143.731 |  |  |  | -9.527 |  | C | 19¦C¦4
19 | Ace Bypass |  | 5 | 0.454 |  |  |  | 1 | 4707 | 88.71 |  |  |  | -25.22 |  | C | 19¦C¦5
19 | Ace Bypass |  | 6 | 0.454 |  |  |  | 0 | 4741 | 88.819 |  |  |  | -25.352 |  | C | 19¦C¦6
19 | Ace Bypass |  | 7 | 0.454 |  |  |  | 0 | 4738 | 88.91 |  |  |  | -25.417 |  | C | 19¦C¦7
18 | Ace Drift3 |  |  | 0.415 | [6 empty] | 1 | 91195 |  |  | 2168
18 | Ace Drift3 |  |  | 0.415 | [6 empty] | 2 | 719426 |  |  | 2772
18 | Ace Drift3 |  | 1 | 0.415 | 658 | 25.8 | 12.735 | 0 |  |  |  |  | -1.264 |  |  | N | 18¦N¦1
18 | Ace Drift3 |  | 2 | 0.415 | 661 | 25.7 | 12.751 | 1 |  |  |  |  | -1.25 |  |  | N | 18¦N¦2
18 | Ace Drift3 |  | 3 | 0.415 | 1611 | 27.3 | 27.655 | 0 |  |  |  |  | -5.945 |  |  | N | 18¦N¦3
18 | Ace Drift3 |  | 4 | 0.415 |  |  |  | 0 | 5794 | 132.985 |  |  |  | -9.453 |  | C | 18¦C¦4
18 | Ace Drift3 |  | 5 | 0.415 |  |  |  | 1 | 4709 | 88.765 |  |  |  | -25.22 |  | C | 18¦C¦5
18 | Ace Drift3 |  | 6 | 0.415 |  |  |  | 0 | 4741 | 88.926 |  |  |  | -25.338 |  | C | 18¦C¦6
18 | Ace Drift3 |  | 7 | 0.415 |  |  |  | 0 | 4726 | 88.819 |  |  |  | -25.387 |  | C | 18¦C¦7
17 | USGS 65 |  |  | 0.423 | [6 empty] | 1 | 121159 |  |  | 2171
17 | USGS 65 |  |  | 0.423 | [6 empty] | 2 | 320834 |  |  | 2761
17 | USGS 65 |  | 1 | 0.423 | 657 | 26 | 12.72 | 0 |  |  |  |  | -1.176 |  |  | N | 17¦N¦1
17 | USGS 65 |  | 2 | 0.423 | 660 | 25.9 | 12.729 | 1 |  |  |  |  | -1.25 |  |  | N | 17¦N¦2
17 | USGS 65 |  | 3 | 0.423 | 2817 | 27.6 | 48.547 | 0 |  |  |  |  | 19.444 |  |  | N | 17¦N¦3
17 | USGS 65 |  | 4 | 0.423 |  |  |  | 0 | 2667 | 58.79 |  |  |  | -3.429 |  | C | 17¦C¦4
17 | USGS 65 |  | 5 | 0.423 |  |  |  | 1 | 4685 | 88.76 |  |  |  | -25.22 |  | C | 17¦C¦5
17 | USGS 65 |  | 6 | 0.423 |  |  |  | 0 | 4729 | 88.879 |  |  |  | -25.359 |  | C | 17¦C¦6
17 | USGS 65 |  | 7 | 0.423 |  |  |  | 0 | 4722 | 88.656 |  |  |  | -25.446 |  | C | 17¦C¦7
16 | USGS 65 |  |  | 0.43 | [6 empty] | 1 | 123818 |  |  | 2173
16 | USGS 65 |  |  | 0.43 | [6 empty] | 2 | 334069 |  |  | 2762
16 | USGS 65 |  | 1 | 0.43 | 655 | 26.1 | 12.684 | 0 |  |  |  |  | -1.125 |  |  | N | 16¦N¦1
16 | USGS 65 |  | 2 | 0.43 | 660 | 26.1 | 12.705 | 1 |  |  |  |  | -1.25 |  |  | N | 16¦N¦2
16 | USGS 65 |  | 3 | 0.43 | 2907 | 27.9 | 50.157 | 0 |  |  |  |  | 19.262 |  |  | N | 16¦N¦3
16 | USGS 65 |  | 4 | 0.43 |  |  |  | 0 | 2740 | 61.156 |  |  |  | -3.422 |  | C | 16¦C¦4
16 | USGS 65 |  | 5 | 0.43 |  |  |  | 1 | 4703 | 88.485 |  |  |  | -25.22 |  | C | 16¦C¦5
16 | USGS 65 |  | 6 | 0.43 |  |  |  | 0 | 4729 | 88.511 |  |  |  | -25.368 |  | C | 16¦C¦6
16 | USGS 65 |  | 7 | 0.43 |  |  |  | 0 | 4725 | 88.8 |  |  |  | -25.444 |  | C | 16¦C¦7
15 | USGS 65 |  |  | 0.402 | [6 empty] | 1 | 119850 |  |  | 2177
15 | USGS 65 |  |  | 0.402 | [6 empty] | 2 | 312631 |  |  | 2767
15 | USGS 65 |  | 1 | 0.402 | 653 | 26.5 | 12.654 | 0 |  |  |  |  | -1.275 |  |  | N | 15¦N¦1
15 | USGS 65 |  | 2 | 0.402 | 659 | 26.4 | 12.677 | 1 |  |  |  |  | -1.25 |  |  | N | 15¦N¦2
15 | USGS 65 |  | 3 | 0.402 | 2745 | 28.2 | 47.392 | 0 |  |  |  |  | 19.505 |  |  | N | 15¦N¦3
15 | USGS 65 |  | 4 | 0.402 |  |  |  | 0 | 2572 | 57.172 |  |  |  | -3.399 |  | C | 15¦C¦4
15 | USGS 65 |  | 5 | 0.402 |  |  |  | 1 | 4687 | 88.565 |  |  |  | -25.22 |  | C | 15¦C¦5
15 | USGS 65 |  | 6 | 0.402 |  |  |  | 0 | 4713 | 88.664 |  |  |  | -25.384 |  | C | 15¦C¦6
15 | USGS 65 |  | 7 | 0.402 |  |  |  | 0 | 4716 | 88.684 |  |  |  | -25.45 |  | C | 15¦C¦7
14 | Ace Drift2 |  |  | 0.428 | [6 empty] | 1 | 92927 |  |  | 2176
14 | Ace Drift2 |  |  | 0.428 | [6 empty] | 2 | 744503 |  |  | 2785
14 | Ace Drift2 |  | 1 | 0.428 | 653 | 26.7 | 12.653 | 0 |  |  |  |  | -1.237 |  |  | N | 14¦N¦1
14 | Ace Drift2 |  | 2 | 0.428 | 657 | 26.7 | 12.658 | 1 |  |  |  |  | -1.25 |  |  | N | 14¦N¦2
14 | Ace Drift2 |  | 3 | 0.428 | 1666 | 28.6 | 28.701 | 0 |  |  |  |  | -5.744 |  |  | N | 14¦N¦3
14 | Ace Drift2 |  | 4 | 0.428 |  |  |  | 0 | 6002 | 136.908 |  |  |  | -9.583 |  | C | 14¦C¦4
14 | Ace Drift2 |  | 5 | 0.428 |  |  |  | 1 | 4703 | 88.498 |  |  |  | -25.22 |  | C | 14¦C¦5
14 | Ace Drift2 |  | 6 | 0.428 |  |  |  | 0 | 4718 | 88.592 |  |  |  | -25.377 |  | C | 14¦C¦6
14 | Ace Drift2 |  | 7 | 0.428 |  |  |  | 0 | 4713 | 88.712 |  |  |  | -25.446 |  | C | 14¦C¦7
13 | USGS 64 |  |  | 0.389 | [6 empty] | 1 | 118673 |  |  | 2178
13 | USGS 64 |  |  | 0.389 | [6 empty] | 2 | 306065 |  |  | 2763
13 | USGS 64 |  | 1 | 0.389 | 651 | 27 | 12.624 | 0 |  |  |  |  | -1.266 |  |  | N | 13¦N¦1
13 | USGS 64 |  | 2 | 0.389 | 656 | 27 | 12.641 | 1 |  |  |  |  | -1.25 |  |  | N | 13¦N¦2
13 | USGS 64 |  | 3 | 0.389 | 2667 | 29 | 46.173 | 0 |  |  |  |  | 0.888 |  |  | N | 13¦N¦3
13 | USGS 64 |  | 4 | 0.389 |  |  |  | 0 | 2543 | 55.493 |  |  |  | -23.656 |  | C | 13¦C¦4
13 | USGS 64 |  | 5 | 0.389 |  |  |  | 1 | 4684 | 88.395 |  |  |  | -25.22 |  | C | 13¦C¦5
13 | USGS 64 |  | 6 | 0.389 |  |  |  | 0 | 4709 | 88.489 |  |  |  | -25.393 |  | C | 13¦C¦6
13 | USGS 64 |  | 7 | 0.389 |  |  |  | 0 | 4726 | 88.502 |  |  |  | -25.489 |  | C | 13¦C¦7
12 | USGS 64 |  |  | 0.409 | [6 empty] | 1 | 120042 |  |  | 2185
12 | USGS 64 |  |  | 0.409 | [6 empty] | 2 | 317157 |  |  | 2797
12 | USGS 64 |  | 1 | 0.409 | 650 | 27.4 | 12.588 | 0 |  |  |  |  | -1.129 |  |  | N | 12¦N¦1
12 | USGS 64 |  | 2 | 0.409 | 655 | 27.3 | 12.595 | 1 |  |  |  |  | -1.25 |  |  | N | 12¦N¦2
12 | USGS 64 |  | 3 | 0.409 | 2800 | 29.5 | 48.488 | 0 |  |  |  |  | 1.272 |  |  | N | 12¦N¦3
12 | USGS 64 |  | 4 | 0.409 |  |  |  | 0 | 2636 | 57.934 |  |  |  | -23.686 |  | C | 12¦C¦4
12 | USGS 64 |  | 5 | 0.409 |  |  |  | 1 | 4705 | 88.348 |  |  |  | -25.22 |  | C | 12¦C¦5
12 | USGS 64 |  | 6 | 0.409 |  |  |  | 0 | 4708 | 88.476 |  |  |  | -25.396 |  | C | 12¦C¦6
12 | USGS 64 |  | 7 | 0.409 |  |  |  | 0 | 4704 | 88.507 |  |  |  | -25.47 |  | C | 12¦C¦7
11 | USGS 64 |  |  | 0.397 | [6 empty] | 1 | 117817 |  |  | 2189
11 | USGS 64 |  |  | 0.397 | [6 empty] | 2 | 308780 |  |  | 2813
11 | USGS 64 |  | 1 | 0.397 | 650 | 27.8 | 12.577 | 0 |  |  |  |  | -1.203 |  |  | N | 11¦N¦1
11 | USGS 64 |  | 2 | 0.397 | 654 | 27.7 | 12.597 | 1 |  |  |  |  | -1.25 |  |  | N | 11¦N¦2
11 | USGS 64 |  | 3 | 0.397 | 2696 | 30 | 46.912 | 0 |  |  |  |  | 0.951 |  |  | N | 11¦N¦3
11 | USGS 64 |  | 4 | 0.397 |  |  |  | 0 | 2564 | 56.217 |  |  |  | -23.653 |  | C | 11¦C¦4
11 | USGS 64 |  | 5 | 0.397 |  |  |  | 1 | 4686 | 88.451 |  |  |  | -25.22 |  | C | 11¦C¦5
11 | USGS 64 |  | 6 | 0.397 |  |  |  | 0 | 4708 | 88.404 |  |  |  | -25.384 |  | C | 11¦C¦6
11 | USGS 64 |  | 7 | 0.397 |  |  |  | 0 | 4710 | 88.364 |  |  |  | -25.442 |  | C | 11¦C¦7
10 | Ace Drift1 |  |  | 0.4 | [6 empty] | 1 | 88131 |  |  | 2195
10 | Ace Drift1 |  |  | 0.4 | [6 empty] | 2 | 686469 |  |  | 2840
10 | Ace Drift1 |  | 1 | 0.4 | 649 | 28.1 | 12.558 | 0 |  |  |  |  | -1.284 |  |  | N | 10¦N¦1
10 | Ace Drift1 |  | 2 | 0.4 | 653 | 28.1 | 12.57 | 1 |  |  |  |  | -1.25 |  |  | N | 10¦N¦2
10 | Ace Drift1 |  | 3 | 0.4 | 1514 | 30.3 | 26.4 | 0 |  |  |  |  | -5.523 |  |  | N | 10¦N¦3
10 | Ace Drift1 |  | 4 | 0.4 |  |  |  | 0 | 5679 | 125.163 |  |  |  | -9.552 |  | C | 10¦C¦4
10 | Ace Drift1 |  | 5 | 0.4 |  |  |  | 1 | 4687 | 88.336 |  |  |  | -25.22 |  | C | 10¦C¦5
10 | Ace Drift1 |  | 6 | 0.4 |  |  |  | 0 | 4711 | 88.411 |  |  |  | -25.36 |  | C | 10¦C¦6
10 | Ace Drift1 |  | 7 | 0.4 |  |  |  | 0 | 4701 | 88.399 |  |  |  | -25.401 |  | C | 10¦C¦7
9 | Ace Lin4 |  |  | 0.788 | [6 empty] | 1 | 120954 |  |  | 2239
9 | Ace Lin4 |  |  | 0.788 | [6 empty] | 2 | 1309408 |  |  | 2878
9 | Ace Lin4 |  | 1 | 0.788 | 649 | 28.6 | 12.573 | 0 |  |  |  |  | -1.323 |  |  | N | 9¦N¦1
9 | Ace Lin4 |  | 2 | 0.788 | 654 | 28.5 | 12.584 | 1 |  |  |  |  | -1.25 |  |  | N | 9¦N¦2
9 | Ace Lin4 |  | 3 | 0.788 | 2995 | 30.8 | 51.069 | 0 |  |  |  |  | -5.662 |  |  | N | 9¦N¦3
9 | Ace Lin4 |  | 4 | 0.788 |  |  |  | 0 | 10285 | 240.154 |  |  |  | -9.352 |  | C | 9¦C¦4
9 | Ace Lin4 |  | 5 | 0.788 |  |  |  | 1 | 4705 | 88.337 |  |  |  | -25.22 |  | C | 9¦C¦5
9 | Ace Lin4 |  | 6 | 0.788 |  |  |  | 0 | 4711 | 88.376 |  |  |  | -25.316 |  | C | 9¦C¦6
9 | Ace Lin4 |  | 7 | 0.788 |  |  |  | 0 | 4686 | 88.438 |  |  |  | -25.351 |  | C | 9¦C¦7
8 | Ace Lin3 |  |  | 0.572 | [6 empty] | 1 | 107156 |  |  | 2232
8 | Ace Lin3 |  |  | 0.572 | [6 empty] | 2 | 974335 |  |  | 2835
8 | Ace Lin3 |  | 1 | 0.572 | 651 | 28.9 | 12.588 | 0 |  |  |  |  | -1.237 |  |  | N | 8¦N¦1
8 | Ace Lin3 |  | 2 | 0.572 | 655 | 28.9 | 12.597 | 1 |  |  |  |  | -1.25 |  |  | N | 8¦N¦2
8 | Ace Lin3 |  | 3 | 0.572 | 2213 | 31.6 | 38.149 | 0 |  |  |  |  | -5.479 |  |  | N | 8¦N¦3
8 | Ace Lin3 |  | 4 | 0.572 |  |  |  | 0 | 7464 | 178.135 |  |  |  | -9.384 |  | C | 8¦C¦4
8 | Ace Lin3 |  | 5 | 0.572 |  |  |  | 1 | 4709 | 88.691 |  |  |  | -25.22 |  | C | 8¦C¦5
8 | Ace Lin3 |  | 6 | 0.572 |  |  |  | 0 | 4720 | 88.61 |  |  |  | -25.396 |  | C | 8¦C¦6
8 | Ace Lin3 |  | 7 | 0.572 |  |  |  | 0 | 4722 | 88.686 |  |  |  | -25.471 |  | C | 8¦C¦7
7 | Ace Lin2 |  |  | 0.177 | [6 empty] | 1 | 70685 |  |  | 2233
7 | Ace Lin2 |  |  | 0.177 | [6 empty] | 2 | 298826 |  |  | 2820
7 | Ace Lin2 |  | 1 | 0.177 | 650 | 29.5 | 12.577 | 0 |  |  |  |  | -1.142 |  |  | N | 7¦N¦1
7 | Ace Lin2 |  | 2 | 0.177 | 656 | 29.4 | 12.599 | 1 |  |  |  |  | -1.25 |  |  | N | 7¦N¦2
7 | Ace Lin2 |  | 3 | 0.177 | 659 | 32.3 | 12.058 | 0 |  |  |  |  | -4.102 |  |  | N | 7¦N¦3
7 | Ace Lin2 |  | 4 | 0.177 |  |  |  | 0 | 2402 | 54.07 |  |  |  | -9.667 |  | C | 7¦C¦4
7 | Ace Lin2 |  | 5 | 0.177 |  |  |  | 1 | 4653 | 88.525 |  |  |  | -25.22 |  | C | 7¦C¦5
7 | Ace Lin2 |  | 6 | 0.177 |  |  |  | 0 | 4698 | 88.62 |  |  |  | -25.457 |  | C | 7¦C¦6
7 | Ace Lin2 |  | 7 | 0.177 |  |  |  | 0 | 4721 | 88.716 |  |  |  | -25.527 |  | C | 7¦C¦7
6 | Ace Lin1 |  |  | 0.126 | [6 empty] | 1 | 67384 |  |  | 2214
6 | Ace Lin1 |  |  | 0.126 | [6 empty] | 2 | 209900 |  |  | 2817
6 | Ace Lin1 |  | 1 | 0.126 | 650 | 30 | 12.579 | 0 |  |  |  |  | -1.207 |  |  | N | 6¦N¦1
6 | Ace Lin1 |  | 2 | 0.126 | 654 | 30 | 12.592 | 1 |  |  |  |  | -1.25 |  |  | N | 6¦N¦2
6 | Ace Lin1 |  | 3 | 0.126 | 466 | 33.1 | 8.889 | 0 |  |  |  |  | -3.024 |  |  | N | 6¦N¦3
6 | Ace Lin1 |  | 4 | 0.126 |  |  |  | 0 | 1684 | 37.895 |  |  |  | -9.666 |  | C | 6¦C¦4
6 | Ace Lin1 |  | 5 | 0.126 |  |  |  | 1 | 4676 | 88.325 |  |  |  | -25.22 |  | C | 6¦C¦5
6 | Ace Lin1 |  | 6 | 0.126 |  |  |  | 0 | 4713 | 88.553 |  |  |  | -25.438 |  | C | 6¦C¦6
6 | Ace Lin1 |  | 7 | 0.126 |  |  |  | 0 | 4737 | 88.662 |  |  |  | -25.505 |  | C | 6¦C¦7
5 | Blank Capsule |  |  | 0 | [6 empty] | 1 | 82725 |  |  | 2216
5 | Blank Capsule |  | 1 | 0 | 650 | 30.8 | 12.577 | 0 |  |  |  |  | -1.071 |  |  | N | 5¦N¦1
5 | Blank Capsule |  | 2 | 0 | 654 | 30.8 | 12.587 | 1 |  |  |  |  | -1.25 |  |  | N | 5¦N¦2
5 | Blank Capsule |  | 3 | 0 |  |  |  | 1 | 4706 | 88.478 |  |  |  | -25.22 |  | C | 5¦C¦3
5 | Blank Capsule |  | 4 | 0 |  |  |  | 0 | 4672 | 88.561 |  |  |  | -25.535 |  | C | 5¦C¦4
5 | Blank Capsule |  | 5 | 0 |  |  |  | 0 | 4686 | 88.619 |  |  |  | -25.635 |  | C | 5¦C¦5
4 | Blank Capsule |  |  | 0 | [6 empty] | 1 | 82276 |  |  | 2241
4 | Blank Capsule |  | 1 | 0 | 650 | 31.9 | 12.574 | 0 |  |  |  |  | -1.133 |  |  | N | 4¦N¦1
4 | Blank Capsule |  | 2 | 0 | 654 | 31.8 | 12.583 | 1 |  |  |  |  | -1.25 |  |  | N | 4¦N¦2
4 | Blank Capsule |  | 3 | 0 | 61 | 35.2 | 1.676 | 0 |  |  |  |  | 15.781 |  |  | N | 4¦N¦3
4 | Blank Capsule |  | 4 | 0 |  |  |  | 1 | 4710 | 88.588 |  |  |  | -25.22 |  | C | 4¦C¦4
4 | Blank Capsule |  | 5 | 0 |  |  |  | 0 | 4672 | 88.654 |  |  |  | -25.478 |  | C | 4¦C¦5
4 | Blank Capsule |  | 6 | 0 |  |  |  | 0 | 4672 | 88.599 |  |  |  | -25.573 |  | C | 4¦C¦6
3 | Ace Bypass |  |  | 0.49 | [6 empty] | 1 | 104175 |  |  | 2233
3 | Ace Bypass |  |  | 0.49 | [6 empty] | 2 | 853203 |  |  | 2830
3 | Ace Bypass |  | 1 | 0.49 | 648 | 32.8 | 12.574 | 0 |  |  |  |  | -1.2 |  |  | N | 3¦N¦1
3 | Ace Bypass |  | 2 | 0.49 | 653 | 32.7 | 12.578 | 1 |  |  |  |  | -1.25 |  |  | N | 3¦N¦2
3 | Ace Bypass |  | 3 | 0.49 | 1952 | 36.4 | 34.829 | 0 |  |  |  |  | -4.515 |  |  | N | 3¦N¦3
3 | Ace Bypass |  | 4 | 0.49 |  |  |  | 0 | 6715 | 155.895 |  |  |  | -9.45 |  | C | 3¦C¦4
3 | Ace Bypass |  | 5 | 0.49 |  |  |  | 1 | 4706 | 88.54 |  |  |  | -25.22 |  | C | 3¦C¦5
3 | Ace Bypass |  | 6 | 0.49 |  |  |  | 0 | 4711 | 88.507 |  |  |  | -25.401 |  | C | 3¦C¦6
3 | Ace Bypass |  | 7 | 0.49 |  |  |  | 0 | 4719 | 88.595 |  |  |  | -25.421 |  | C | 3¦C¦7
2 | Ace Bypass |  |  | 1.131 | [6 empty] | 1 | 163029 |  |  | 2217
2 | Ace Bypass |  |  | 1.131 | [6 empty] | 2 | 1962837 |  |  | 2885
2 | Ace Bypass |  | 1 | 1.131 | 651 | 33.3 | 12.572 | 0 |  |  |  |  | -1.139 |  |  | N | 2¦N¦1
2 | Ace Bypass |  | 2 | 1.131 | 654 | 33.2 | 12.585 | 1 |  |  |  |  | -1.25 |  |  | N | 2¦N¦2
2 | Ace Bypass |  | 3 | 1.131 | 4689 | 37.2 | 79.936 | 0 |  |  |  |  | -4.989 |  |  | N | 2¦N¦3
2 | Ace Bypass |  | 4 | 1.131 |  |  |  | 0 | 14212 | 362.648 |  |  |  | -9.107 |  | C | 2¦C¦4
2 | Ace Bypass |  | 5 | 1.131 |  |  |  | 1 | 4715 | 88.485 |  |  |  | -25.22 |  | C | 2¦C¦5
2 | Ace Bypass |  | 6 | 1.131 |  |  |  | 0 | 4733 | 88.757 |  |  |  | -25.353 |  | C | 2¦C¦6
2 | Ace Bypass |  | 7 | 1.131 |  |  |  | 0 | 4692 | 88.545 |  |  |  | -25.43 |  | C | 2¦C¦7
1 | Blank Hole |  |  | 0 | [6 empty] | 1 | 86729 |  |  | 2218
1 | Blank Hole |  | 1 | 0 | 651 | 34.3 | 12.585 | 0 |  |  |  |  | -1.217 |  |  | N | 1¦N¦1
1 | Blank Hole |  | 2 | 0 | 655 | 34.2 | 12.597 | 1 |  |  |  |  | -1.25 |  |  | N | 1¦N¦2
1 | Blank Hole |  | 3 | 0 | 217 | 38.6 | 6.018 | 0 |  |  |  |  | 3.728 |  |  | N | 1¦N¦3
1 | Blank Hole |  | 4 | 0 |  |  |  | 1 | 4705 | 88.098 |  |  |  | -25.22 |  | C | 1¦C¦4
1 | Blank Hole |  | 5 | 0 |  |  |  | 0 | 4677 | 88.349 |  |  |  | -25.736 |  | C | 1¦C¦5
1 | Blank Hole |  | 6 | 0 |  |  |  | 0 | 4677 | 88.576 |  |  |  | -25.918 |  | C | 1¦C¦6

## Setup
Setup
Yellow cells are yours. Everything else on this sheet is a label.
Run information
Run name | 20260915 TestRun
Run date | 2026-09-15
Analyst
Processing options
Blank correction | Auto |  |  |  |  | On / Off / Auto. Auto applies it only when at least three blanks agree to within 2 permil.
Linearity correction | Auto |  |  |  |  | On / Off / Auto. Auto applies the fitted slope only when it differs from zero at the significance level below.
Drift correction | Auto |  |  |  |  | On / Off / Auto. Auto applies the fitted slope only when it differs from zero at the significance level below.
Drift model | Row |  |  |  |  | Row / Background / Auto. Auto takes the lower residual scatter, which is not an independent test. See the note under the table.
First position to use in the drift fit | [5 empty] | Leave empty to use every position. QC suggests a value. The calibration anchors are never excluded, since they are always run early.
Significance level for Auto | 0.05 |  |  |  |  | A slope is applied on Auto when its two-sided p-value is below this. The FIRMS guide asks that a correction be shown necessary before it is applied.
Positions left out of the nitrogen drift fit | [5 empty] | Comma-separated autosampler positions, e.g. 59. Use it for a drift monitor QC calls an outlier when the cause is the capsule, not the instrument.
Positions left out of the carbon drift fit | [5 empty] | Separate lists, because a capsule can go wrong in one element only.
Flag thresholds
Check standard offset (permil) | 0.5 |  |  |  |  | Flag a check standard further than this from its accepted value.
Minimum Area 28 | [5 empty] | Leave empty to use the smallest linearity standard in the run.
Minimum Area 44 | [5 empty] | Leave empty to use the smallest linearity standard in the run.
Maximum Area 28 | [5 empty] | Leave empty to use the largest linearity standard in the run.
Maximum Area 44 | [5 empty] | Leave empty to use the largest linearity standard in the run.
Minimum peak area, x the blank | 3 |  |  |  |  | Below this multiple of the blank the blank correction amplifies rather than corrects. Leave empty to skip the check.
Settled when background excess is under | 0.1 |  |  |  |  | Fraction of the run's starting excess left in the reference-peak background. Used only to suggest a first position and to flag.
Drift monitor outlier, studentized residual | 4 |  |  |  |  | A drift monitor further than this from the line through the others, in units of their scatter, is flagged.
Carryover: background rise after an oversized sample (mV) | 2 |  |  |  |  | Flags a position when one of the two before it is above the linearity range and its sample-peak background rose by more than this. Leave empty to skip the check.
Uncertainty
Sample repeatability, d15N (1 SD, permil) | [5 empty] | One analysis of one sample. Leave empty to use the pooled scatter of the anchor replicates in this run; better is the long-term SD of a matrix-matched in-house standard.
Sample repeatability, d13C (1 SD, permil)
Coverage factor k | 2 |  |  |  |  | Expanded uncertainty U = k x u. k = 2 is about 95 percent only when the replicates behind u are many; with a handful, take k from Student's t.
Standards
Accepted values below are the published USGS values as of 2026-09-16, and u is the 1-sigma uncertainty the producer lists. Check both against your certificates.
Match text in Identifier 1 | Role | Accepted d13C | Accepted d15N | %C | %N | Material | u(d13C) | u(d15N)
USGS 64 | Anchor | -40.81 | 1.76 | 32 | 18.66 | glycine | 0.04 | 0.06
USGS 65 | Anchor | -20.29 | 20.68 | 32 | 18.66 | glycine | 0.04 | 0.06
USGS 66 | Check | -0.67 | 40.83 | 32 | 18.66 | glycine | 0.04 | 0.06
USGS 40 | Check | -26.39 | -4.52 | 40.82 | 9.52 | L-glutamic acid
USGS 41a | Check | 36.55 | 47.55 | 40.82 | 9.52 | L-glutamic acid
Ace Lin | Linearity |  |  | 71.09 | 10.36 | acetanilide, size series
Ace Drift | Drift |  |  | 71.09 | 10.36 | acetanilide, drift monitor
Ace Bypass | Exclude |  |  | 71.09 | 10.36 | bypass test
Blank Capsule | Blank |  |  |  |  | tin capsule, nothing in it
Blank Hole | Exclude |  |  |  |  | empty carousel position
Constants
Atomic mass C (g/mol) | 12.011 |  |  |  |  | IUPAC
Atomic mass N (g/mol) | 14.007 |  |  |  |  | IUPAC

## Work
Work
One row per autosampler position. Type sample masses in the yellow column F, in mg. Everything else is a formula. Grey columns are machinery.
identification |  |  |  | mass |  |  | peak lookup |  |  |  | as measured | [6 empty] | from the standards table |  |  |  | nitrogen corrections |  |  |  |  | carbon corrections |  |  |  |  | elemental content | [5 empty] | fit inputs and helpers, ignore | [21 empty] | instrument settling |  | fit inputs and helpers, ignore | [7 empty] | reference pulses |  |  |  |  | fit inputs and helpers, ignore | [13 empty] | uncertainty, 1 SD and k x SD |  |  |  | fit inputs and helpers, ignore | [15 empty] | chart feeds, ignore | [28 empty] | carryover and settling helpers, ignore
Row | Identifier | Std row | Type | Mass from export (mg) | Mass entered (mg) | Mass used (mg) | peak N | key N | peak C | key C | Area 28 | Ampl 28 | BGD 28 | d15N raw | Area 44 | Ampl 44 | d13C raw | Accepted d15N | Accepted d13C | Certified %N | Certified %C | d15N blank-corr | d15N linearity-corr | drift x (N) | d15N drift-corr | d15N calibrated | d13C blank-corr | d13C linearity-corr | drift x (C) | d13C drift-corr | d13C calibrated | ug N | ug C | %N | %C | C:N molar | C:N mass | lin x N | lin y N | lin x C | lin y C | drift x row | drift x bgd | drift y N | drift y C | cal x N | cal y N | cal x C | cal y C | kfit x N | kfit y N | kfit x C | kfit y C | blank d15N | blank A28 | blank d13C | blank A44 | drift cal d15N | drift cal d13C | ref BGD 28 | settling frac | drift x early | drift y early | drift x late | drift y late | stage raw | stage blank | stage lin | stage drift | ref peak N | ref pulse d15N | ref flagged C | ref peak C | ref pulse d13C | excl N | excl C | drift x N | drift x C | t drift N | t drift C | cal one N | cal dev N | cal res N | cal sq N | cal one C | cal dev C | cal res C | cal sq C | u d15N | U d15N | u d13C | U d13C | settle streak | ref x N | ref x C | match 1 | match 2 | match 3 | match 4 | match 5 | match 6 | match 7 | match 8 | match 9 | match 10 | match 11 | match 12 | match 13 | ch pos | ch bgd | ch bgd sample | ch bgd monitor | mon bgd | mon bgd x | ch lin x N | ch lin y N | ch samp x N | ch samp y N | ch drift used N | ch drift out N | ch drift after N | ch cal y N | ch acc min N | ch acc max N | ch pulse N | ch lin x C | ch lin y C | ch samp x C | ch samp y C | ch drift used C | ch drift out C | ch drift after C | ch cal y C | ch acc min C | ch acc max C | ch pulse C | ch carry | pos used | oversized | exposed | carry ref | carry rise | carry flag | carry rank | after oversized | ref first5 | settle start
1 | Blank Hole | 46 | Exclude |  |  |  | 3 | 1¦N¦3 | 4 | 1¦C¦4 | 6.018 | 217 | 38.6 | 3.728 | [7 empty] | 3.728 | 3.728 | 1 | 3.2445 | 4.3581 | [5 empty] | 6.6624 | [27 empty] | 34.2 | 1 | [8 empty] | 1 | -1.217 | 4 | 5 | -25.736 | 0 | 0 | [12 empty] | 0.20925 | 0.41849 |  |  | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 46 | 0 | 0 | 0 | 1 | 34.2 | 38.6 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.217 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.736 | #N/A | 1 | 0 | 0 | 38.6 |  | 0 |  | 0 | 34.2
2 | Ace Bypass | 44 | Exclude | 1.131 |  | 1.131 | 3 | 2¦N¦3 | 4 | 2¦C¦4 | 79.936 | 4689 | 37.2 | -4.989 | 362.648 | 14212 | -9.107 |  |  | 10.36 | 71.09 | -4.989 | -4.989 | 2 | -5.4577 | -4.5756 | -9.107 | -9.5143 | 2 | -9.4318 | -26.5575 | 128.7681 | 822.0055 | 11.3853 | 72.6795 | 7.4444 | 6.3836 | [22 empty] | 33.2 | 0.90991 | [8 empty] | 1 | -1.139 | 5 | 6 | -25.353 | 0 | 0 | [12 empty] | 0.25209 | 0.50419 | 0.03793 | 0.07586 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 44 | 0 | 0 | 0 | 0 | 0 | 2 | 33.2 | 37.2 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.139 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.353 | #N/A | 2 | 1 | 0 | 37.2 |  | 0 |  | 0 | 33.2
3 | Ace Bypass | 44 | Exclude | 0.49 |  | 0.49 | 3 | 3¦N¦3 | 4 | 3¦C¦4 | 34.829 | 1952 | 36.4 | -4.515 | 155.895 | 6715 | -9.45 |  |  | 10.36 | 71.09 | -4.515 | -4.515 | 3 | -4.969 | -4.0739 | -9.45 | -9.4991 | 3 | -9.4192 | -26.5447 | 54.2555 | 351.0937 | 11.0726 | 71.6518 | 7.5465 | 6.4711 | [22 empty] | 32.7 | 0.86486 | [8 empty] | 1 | -1.2 | 5 | 6 | -25.401 | 0 | 0 | [12 empty] | 0.24924 | 0.49848 | 0.037941 | 0.075883 | 0 | 3 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 44 | 0 | 0 | 0 | 0 | 0 | 3 | 32.7 | 36.4 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.2 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.401 | #N/A | 3 | 0 | 1 | 37.2 | -0.8 | 0 |  | 0 | 32.7
4 | Blank Capsule | 45 | Blank |  |  |  | 3 | 4¦N¦3 | 4 | 4¦C¦4 | 1.676 | 61 | 35.2 | 15.781 | [7 empty] | 15.781 | 15.781 | 4 | 15.3418 | 16.7771 | [5 empty] | -0.51022 | [21 empty] | 15.781 | 1.676 |  |  |  |  | 31.8 | 0.78378 | [8 empty] | 1 | -1.133 | 4 | 5 | -25.478 | 0 | 0 | [12 empty] | 0.19096 | 0.38192 |  |  | 0 | 4 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 45 | 0 | 0 | 0 | 0 | 4 | 31.8 | 35.2 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.133 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.478 | #N/A | 4 | 0 | 1 | 37.2 | -2 | 0 |  | 0 | 31.8
5 | Blank Capsule | 45 | Blank |  |  |  | 2 | 5¦N¦2 | 3 | 5¦C¦3 | [49 empty] | 30.8 | 0.69369 | [8 empty] | 1 | -1.071 | 3 | 4 | -25.535 | 0 | 0 | [16 empty] | 0 | 5 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 45 | 0 | 0 | 0 | 0 | 5 | 30.8 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.071 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.535 | #N/A | 5 | 0 | 0 | 37.2 |  | 0 |  | 0 | 30.8
6 | Ace Lin1 | 42 | Linearity | 0.126 |  | 0.126 | 3 | 6¦N¦3 | 4 | 6¦C¦4 | 8.889 | 466 | 33.1 | -3.024 | 37.895 | 1684 | -9.666 |  |  | 10.36 | 71.09 | -3.024 | -3.024 | 6 | -3.4337 | -2.4978 | -9.666 | -9.5106 | 6 | -9.4385 | -26.5643 | 11.405 | 82.3306 | 9.0516 | 65.3417 | 8.4184 | 7.2188 | 8.889 | -3.024 | 37.895 | -9.666 | [8 empty] | 13.0536 | 8.889 | 89.5734 | 37.895 | [6 empty] | 30 | 0.62162 | [8 empty] | 1 | -1.207 | 5 | 6 | -25.438 | 0 | 0 | [12 empty] | 0.24056 | 0.48113 | 0.037924 | 0.075849 | 0 | 6 | 6 | 0 | 0 | 0 | 0 | 0 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 30 | 33.1 | #N/A |  |  | 8.889 | -3.024 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.207 | 37.895 | -9.666 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.438 | #N/A | 6 | 0 | 0 | 33.1 |  | 0 |  | 0
7 | Ace Lin2 | 42 | Linearity | 0.177 |  | 0.177 | 3 | 7¦N¦3 | 4 | 7¦C¦4 | 12.058 | 659 | 32.3 | -4.102 | 54.07 | 2402 | -9.667 |  |  | 10.36 | 71.09 | -4.102 | -4.102 | 7 | -4.497 | -3.5893 | -9.667 | -9.5397 | 7 | -9.4701 | -26.5964 | 16.6399 | 119.1716 | 9.4011 | 67.3286 | 8.352 | 7.1618 | 12.058 | -4.102 | 54.07 | -9.667 | [8 empty] | 18.3372 | 12.058 | 125.8293 | 54.07 | [6 empty] | 29.4 | 0.56757 | [8 empty] | 1 | -1.142 | 5 | 6 | -25.457 | 0 | 0 | [12 empty] | 0.24652 | 0.49304 | 0.037897 | 0.075793 | 0 | 7 | 7 | 0 | 0 | 0 | 0 | 0 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 29.4 | 32.3 | #N/A |  |  | 12.058 | -4.102 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.142 | 54.07 | -9.667 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.457 | #N/A | 7 | 0 | 0 | 32.3 |  | 0 |  | 0
8 | Ace Lin3 | 42 | Linearity | 0.572 |  | 0.572 | 3 | 8¦N¦3 | 4 | 8¦C¦4 | 38.149 | 2213 | 31.6 | -5.479 | 178.135 | 7464 | -9.384 |  |  | 10.36 | 71.09 | -5.479 | -5.479 | 8 | -5.8592 | -4.9878 | -9.384 | -9.4716 | 8 | -9.4047 | -26.53 | 59.7398 | 401.7487 | 10.444 | 70.2358 | 7.8425 | 6.725 | 38.149 | -5.479 | 178.135 | -9.384 | [8 empty] | 59.2592 | 38.149 | 406.6348 | 178.135 | [6 empty] | 28.9 | 0.52252 | [8 empty] | 1 | -1.237 | 5 | 6 | -25.396 | 0 | 0 | [12 empty] | 0.25447 | 0.50894 | 0.037954 | 0.075909 | 0 | 8 | 8 | 0 | 0 | 0 | 0 | 0 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 28.9 | 31.6 | #N/A |  |  | 38.149 | -5.479 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.237 | 178.135 | -9.384 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.396 | #N/A | 8 | 0 | 0 | 31.6 |  | 0 |  | 0
9 | Ace Lin4 | 42 | Linearity | 0.788 |  | 0.788 | 3 | 9¦N¦3 | 4 | 9¦C¦4 | 51.069 | 2995 | 30.8 | -5.662 | 240.154 | 10285 | -9.352 |  |  | 10.36 | 71.09 | -5.662 | -5.662 | 9 | -6.0275 | -5.1605 | -9.352 | -9.5471 | 9 | -9.4827 | -26.6092 | 81.0825 | 543.0065 | 10.2897 | 68.9095 | 7.8099 | 6.697 | 51.069 | -5.662 | 240.154 | -9.352 | [8 empty] | 81.6368 | 51.069 | 560.1892 | 240.154 | [6 empty] | 28.5 | 0.48649 | [8 empty] | 1 | -1.323 | 5 | 6 | -25.316 | 0 | 0 | [12 empty] | 0.25548 | 0.51095 | 0.037885 | 0.075771 | 0 | 9 | 9 | 0 | 0 | 0 | 0 | 0 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9 | 28.5 | 30.8 | #N/A |  |  | 51.069 | -5.662 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.323 | 240.154 | -9.352 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.316 | #N/A | 9 | 0 | 0 | 30.8 |  | 0 |  | 0
10 | Ace Drift1 | 43 | Drift | 0.4 |  | 0.4 | 3 | 10¦N¦3 | 4 | 10¦C¦4 | 26.4 | 1514 | 30.3 | -5.523 | 125.163 | 5679 | -9.552 |  |  | 10.36 | 71.09 | -5.523 | -5.523 | 10 | -5.8737 | -5.0027 | -9.552 | -9.5478 | 10 | -9.4861 | -26.6126 | 40.3316 | 281.0969 | 10.0829 | 70.2742 | 8.1279 | 6.9696 |  |  |  |  | 10 | 28.1 | -5.523 | -9.5478 |  |  |  |  | 41.44 | 26.4 | 284.36 | 125.163 |  |  |  |  | -5.0027 | -26.6126 | 28.1 | 0.45045 | 10 | -5.523 |  |  | -5.523 | -5.523 | -5.523 | -5.8737 | 1 | -1.284 | 5 | 6 | -25.36 | 0 | 0 | 10 | 10 | 2.1951 | -0.1785 | [8 empty] | 0.25456 | 0.50912 | 0.037883 | 0.075765 | 0 | 10 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 28.1 | 30.3 | 30.3 | 30.3 | 10 | #N/A | #N/A | #N/A | #N/A | -5.523 | #N/A | -5.8737 | #N/A | 1.76 | 20.68 | -1.284 | #N/A | #N/A | #N/A | #N/A | -9.5478 | #N/A | -9.4861 | #N/A | -40.81 | -20.29 | -25.36 | #N/A | 10 | 0 | 0 | 30.3 |  | 0 |  | 0
11 | USGS 64 | 37 | Anchor | 0.397 |  | 0.397 | 3 | 11¦N¦3 | 4 | 11¦C¦4 | 46.912 | 2696 | 30 | 0.951 | 56.217 | 2564 | -23.653 | 1.76 | -40.81 | 18.66 | 32 | 0.951 | 0.951 | 11 | 0.61505 | 1.6587 | -23.653 | -23.5294 | 11 | -23.4702 | -40.7946 | 74.2155 | 124.0617 | 18.6941 | 31.2498 | 1.9494 | 1.6716 | [8 empty] | 0.61505 | 1.76 | -23.4702 | -40.81 | 74.0802 | 46.912 | 127.04 | 56.217 | [6 empty] | 27.7 | 0.41441 | [8 empty] | 1 | -1.203 | 5 | 6 | -25.384 | 0 | 0 |  |  |  |  | 1 | -9.3136 | 9.6626 | 0.37829 | 1 | -10.1016 | 10.2292 | 550.8516 | 0.22018 | 0.44035 | 0.046654 | 0.093308 | 0 | 11 | 11 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 11 | 27.7 | 30 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.203 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.384 | #N/A | 11 | 0 | 0 | 30 |  | 0 |  | 0
12 | USGS 64 | 37 | Anchor | 0.409 |  | 0.409 | 3 | 12¦N¦3 | 4 | 12¦C¦4 | 48.488 | 2800 | 29.5 | 1.272 | 57.934 | 2636 | -23.686 | 1.76 | -40.81 | 18.66 | 32 | 1.272 | 1.272 | 12 | 0.95081 | 2.0034 | -23.686 | -23.5653 | 12 | -23.5088 | -40.8337 | 76.8189 | 127.9725 | 18.7821 | 31.2891 | 1.9427 | 1.6659 | [8 empty] | 0.95081 | 1.76 | -23.5088 | -40.81 | 76.3194 | 48.488 | 130.88 | 57.934 | [6 empty] | 27.3 | 0.37838 | [8 empty] | 1 | -1.129 | 5 | 6 | -25.396 | 0 | 0 |  |  |  |  | 1 | -8.9778 | 8.9733 | 0.90404 | 1 | -10.1402 | 10.3074 | 552.6637 | 0.21867 | 0.43734 | 0.046725 | 0.093451 | 0 | 12 | 12 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 27.3 | 29.5 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.129 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.396 | #N/A | 12 | 0 | 0 | 29.5 |  | 0 |  | 0
13 | USGS 64 | 37 | Anchor | 0.389 |  | 0.389 | 3 | 13¦N¦3 | 4 | 13¦C¦4 | 46.173 | 2667 | 29 | 0.888 | 55.493 | 2543 | -23.656 | 1.76 | -40.81 | 18.66 | 32 | 0.888 | 0.888 | 13 | 0.58156 | 1.6243 | -23.656 | -23.5311 | 13 | -23.4772 | -40.8016 | 72.9947 | 122.4127 | 18.7647 | 31.4686 | 1.9557 | 1.677 | [8 empty] | 0.58156 | 1.76 | -23.4772 | -40.81 | 72.5874 | 46.173 | 124.48 | 55.493 | [6 empty] | 27 | 0.35135 | [8 empty] | 1 | -1.266 | 5 | 6 | -25.393 | 0 | 0 |  |  |  |  | 1 | -9.3471 | 9.7314 | 0.33822 | 1 | -10.1086 | 10.2432 | 551.1774 | 0.22033 | 0.44066 | 0.046667 | 0.093334 | 0 | 13 | 13 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13 | 27 | 29 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.266 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.393 | #N/A | 13 | 0 | 0 | 29 |  | 0 |  | 0
14 | Ace Drift2 | 43 | Drift | 0.428 |  | 0.428 | 3 | 14¦N¦3 | 4 | 14¦C¦4 | 28.701 | 1666 | 28.6 | -5.744 | 136.908 | 6002 | -9.583 |  |  | 10.36 | 71.09 | -5.744 | -5.744 | 14 | -6.0357 | -5.1689 | -9.583 | -9.5992 | 14 | -9.5478 | -26.6752 | 44.1326 | 307.8479 | 10.3114 | 71.9271 | 8.1347 | 6.9755 |  |  |  |  | 14 | 26.7 | -5.744 | -9.5992 |  |  |  |  | 44.3408 | 28.701 | 304.2652 | 136.908 |  |  |  |  | -5.1689 | -26.6752 | 26.7 | 0.32432 | 14 | -5.744 |  |  | -5.744 | -5.744 | -5.744 | -6.0357 | 1 | -1.237 | 5 | 6 | -25.377 | 0 | 0 | 14 | 14 | 0.51184 | -1.937 | [8 empty] | 0.25553 | 0.51105 | 0.037829 | 0.075658 | 0 | 14 | 14 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 14 | 26.7 | 28.6 | 28.6 | 28.6 | 14 | #N/A | #N/A | #N/A | #N/A | -5.744 | #N/A | -6.0357 | #N/A | 1.76 | 20.68 | -1.237 | #N/A | #N/A | #N/A | #N/A | -9.5992 | #N/A | -9.5478 | #N/A | -40.81 | -20.29 | -25.377 | #N/A | 14 | 0 | 0 | 28.6 |  | 0 |  | 0
15 | USGS 65 | 38 | Anchor | 0.402 |  | 0.402 | 3 | 15¦N¦3 | 4 | 15¦C¦4 | 47.392 | 2745 | 28.2 | 19.505 | 57.172 | 2572 | -3.399 | 20.68 | -20.29 | 18.66 | 32 | 19.505 | 19.505 | 15 | 19.2281 | 20.7668 | -3.399 | -3.277 | 15 | -3.2283 | -20.2662 | 75.0084 | 126.2369 | 18.6588 | 31.4022 | 1.9626 | 1.683 | [8 empty] | 19.2281 | 20.68 | -3.2283 | -20.29 | 75.0132 | 47.392 | 128.64 | 57.172 | [6 empty] | 26.4 | 0.2973 | [8 empty] | 1 | -1.275 | 5 | 6 | -25.384 | 0 | 0 |  |  |  |  | 1 | 9.2994 | -9.6335 | 369.7188 | 1 | 10.1403 | -10.3076 | 10.4217 | 0.19772 | 0.39544 | 0.046722 | 0.093444 | 0 | 15 | 15 | 0 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 26.4 | 28.2 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.275 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.384 | #N/A | 15 | 0 | 0 | 28.2 |  | 0 |  | 0
16 | USGS 65 | 38 | Anchor | 0.43 |  | 0.43 | 3 | 16¦N¦3 | 4 | 16¦C¦4 | 50.157 | 2907 | 27.9 | 19.262 | 61.156 | 2740 | -3.422 | 20.68 | -20.29 | 18.66 | 32 | 19.262 | 19.262 | 16 | 18.9998 | 20.5324 | -3.422 | -3.3069 | 16 | -3.2608 | -20.2992 | 79.576 | 135.3111 | 18.506 | 31.4677 | 1.983 | 1.7004 | [8 empty] | 18.9998 | 20.68 | -3.2608 | -20.29 | 80.238 | 50.157 | 137.6 | 61.156 | [6 empty] | 26.1 | 0.27027 | [8 empty] | 1 | -1.125 | 5 | 6 | -25.368 | 0 | 0 |  |  |  |  | 1 | 9.0712 | -9.1649 | 360.9934 | 1 | 10.1078 | -10.2417 | 10.6326 | 0.19716 | 0.39432 | 0.046662 | 0.093324 | 0 | 16 | 16 | 0 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 26.1 | 27.9 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.125 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.368 | #N/A | 16 | 0 | 0 | 27.9 |  | 0 |  | 0
17 | USGS 65 | 38 | Anchor | 0.423 |  | 0.423 | 3 | 17¦N¦3 | 4 | 17¦C¦4 | 48.547 | 2817 | 27.6 | 19.444 | 58.79 | 2667 | -3.429 | 20.68 | -20.29 | 18.66 | 32 | 19.444 | 19.444 | 17 | 19.1966 | 20.7344 | -3.429 | -3.3098 | 17 | -3.2663 | -20.3047 | 76.9164 | 129.9221 | 18.1835 | 30.7145 | 1.9698 | 1.6891 | [8 empty] | 19.1966 | 20.68 | -3.2663 | -20.29 | 78.9318 | 48.547 | 135.36 | 58.79 | [6 empty] | 25.9 | 0.25225 | [8 empty] | 1 | -1.176 | 5 | 6 | -25.359 | 0 | 0 |  |  |  |  | 1 | 9.2679 | -9.5689 | 368.5087 | 1 | 10.1023 | -10.2305 | 10.6685 | 0.19764 | 0.39528 | 0.046652 | 0.093303 | 0 | 17 | 17 | 0 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 25.9 | 27.6 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.176 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.359 | #N/A | 17 | 0 | 0 | 27.6 |  | 0 |  | 0
18 | Ace Drift3 | 43 | Drift | 0.415 |  | 0.415 | 3 | 18¦N¦3 | 4 | 18¦C¦4 | 27.655 | 1611 | 27.3 | -5.945 | 132.985 | 5794 | -9.453 |  |  | 10.36 | 71.09 | -5.945 | -5.945 | 18 | -6.1777 | -5.3147 | -9.453 | -9.4624 | 18 | -9.4214 | -26.547 | 42.4047 | 298.9127 | 10.218 | 72.0271 | 8.2205 | 7.049 |  |  |  |  | 18 | 25.7 | -5.945 | -9.4624 |  |  |  |  | 42.994 | 27.655 | 295.0235 | 132.985 |  |  |  |  | -5.3147 | -26.547 | 25.7 | 0.23423 | 18 | -5.945 |  |  | -5.945 | -5.945 | -5.945 | -6.1777 | 1 | -1.264 | 5 | 6 | -25.338 | 0 | 0 | 18 | 18 | -0.55857 | 1.4753 | [8 empty] | 0.25638 | 0.51276 | 0.037939 | 0.075879 | 0 | 18 | 18 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 25.7 | 27.3 | 27.3 | 27.3 | 18 | #N/A | #N/A | #N/A | #N/A | -5.945 | #N/A | -6.1777 | #N/A | 1.76 | 20.68 | -1.264 | #N/A | #N/A | #N/A | #N/A | -9.4624 | #N/A | -9.4214 | #N/A | -40.81 | -20.29 | -25.338 | #N/A | 18 | 0 | 0 | 27.3 |  | 0 |  | 0
19 | Ace Bypass | 44 | Exclude | 0.454 |  | 0.454 | 3 | 19¦N¦3 | 4 | 19¦C¦4 | 29.846 | 1745 | 26.9 | -5.941 | 143.731 | 6256 | -9.527 |  |  | 10.36 | 71.09 | -5.941 | -5.941 | 19 | -6.1589 | -5.2955 | -9.527 | -9.555 | 19 | -9.5166 | -26.6436 | 46.024 | 323.3883 | 10.1375 | 71.2309 | 8.1942 | 7.0265 | [22 empty] | 25.4 | 0.20721 | [8 empty] | 1 | -1.131 | 5 | 6 | -25.352 | 0 | 0 | [12 empty] | 0.25627 | 0.51253 | 0.037856 | 0.075712 | 0 | 19 | 19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 44 | 0 | 0 | 0 | 0 | 0 | 19 | 25.4 | 26.9 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.131 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.352 | #N/A | 19 | 0 | 0 | 26.9 |  | 0 |  | 0
20 | Random | 0 | Sample |  |  |  | 3 | 20¦N¦3 | 4 | 20¦C¦4 | 29.179 | 1705 | 26.7 | -5.819 | 139.977 | 6159 | -9.593 |  |  |  |  | -5.819 | -5.819 | 20 | -6.0222 | -5.1551 | -9.593 | -9.6145 | 20 | -9.5787 | -26.7065 | 44.9222 | 314.838 |  |  | 8.1732 | 7.0085 | [22 empty] | 25.3 | 0.1982 | [8 empty] | 1 | -1.305 | 5 | 6 | -25.373 | 0 | 0 | [12 empty] | 0.25545 | 0.51089 | 0.037803 | 0.075605 | 0 | 20 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 25.3 | 26.7 | #N/A |  |  | #N/A | #N/A | 29.179 | -6.0577 | #N/A | #N/A | #N/A | -5.1551 | 1.76 | 20.68 | -1.305 | #N/A | #N/A | 139.977 | -9.717 | #N/A | #N/A | #N/A | -26.7065 | -40.81 | -20.29 | -25.373 | #N/A | 20 | 0 | 0 | 26.7 |  | 0 |  | 0
21 | Random | 0 | Sample |  |  |  | 3 | 21¦N¦3 | 4 | 21¦C¦4 | 10.11 | 577 | 26.4 | -5.969 | 49.134 | 2216 | -9.66 |  |  |  |  | -5.969 | -5.969 | 21 | -6.1574 | -5.2939 | -9.66 | -9.5241 | 21 | -9.4909 | -26.6175 | 13.422 | 107.9291 |  |  | 9.3775 | 8.0412 | [22 empty] | 25.1 | 0.18018 | [8 empty] | 1 | -1.121 | 5 | 6 | -25.441 | 0 | 0 | [12 empty] | 0.25626 | 0.51251 | 0.037878 | 0.075757 | 0 | 21 | 21 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 21 | 25.1 | 26.4 | #N/A |  |  | #N/A | #N/A | 10.11 | -6.0577 | #N/A | #N/A | #N/A | -5.2939 | 1.76 | 20.68 | -1.121 | #N/A | #N/A | 49.134 | -9.717 | #N/A | #N/A | #N/A | -26.6175 | -40.81 | -20.29 | -25.441 | #N/A | 21 | 0 | 0 | 26.4 |  | 0 |  | 0
22 | Ace Drift4 | 43 | Drift | 0.402 |  | 0.402 | 3 | 22¦N¦3 | 4 | 22¦C¦4 | 27.288 | 1591 | 26.1 | -5.874 | 130.484 | 5751 | -9.538 |  |  | 10.36 | 71.09 | -5.874 | -5.874 | 22 | -6.0476 | -5.1812 | -9.538 | -9.5431 | 22 | -9.5125 | -26.6393 | 41.7985 | 293.2163 | 10.3976 | 72.9394 | 8.1808 | 7.015 |  |  |  |  | 22 | 24.8 | -5.874 | -9.5431 |  |  |  |  | 41.6472 | 27.288 | 285.7818 | 130.484 |  |  |  |  | -5.1812 | -26.6393 | 24.8 | 0.15315 | 22 | -5.874 |  |  | -5.874 | -5.874 | -5.874 | -6.0476 | 1 | -1.244 | 5 | 6 | -25.363 | 0 | 0 | 22 | 22 | 0.39572 | -0.78396 | [8 empty] | 0.2556 | 0.5112 | 0.03786 | 0.075719 | 0 | 22 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 22 | 24.8 | 26.1 | 26.1 | 26.1 | 22 | #N/A | #N/A | #N/A | #N/A | -5.874 | #N/A | -6.0476 | #N/A | 1.76 | 20.68 | -1.244 | #N/A | #N/A | #N/A | #N/A | -9.5431 | #N/A | -9.5125 | #N/A | -40.81 | -20.29 | -25.363 | #N/A | 22 | 0 | 0 | 26.1 |  | 0 |  | 0
23 | Random | 0 | Sample |  |  |  | 3 | 23¦N¦3 | 4 | 23¦C¦4 | 59.328 | 3433 | 26 | -3.251 | 18.131 | 825 | -18.113 |  |  |  |  | -3.251 | -3.251 | 23 | -3.4099 | -2.4733 | -18.113 | -17.9234 | 23 | -17.8954 | -35.1409 | 94.7256 | 37.315 |  |  | 0.45939 | 0.39393 | [22 empty] | 24.7 | 0.14414 | [8 empty] | 1 | -1.259 | 5 | 6 | -25.458 | 0 | 0 | [12 empty] | 0.24043 | 0.48087 | 0.038489 | 0.076978 | 0 | 23 | 23 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 23 | 24.7 | 26 | #N/A |  |  | #N/A | #N/A | 59.328 | -6.0577 | #N/A | #N/A | #N/A | -2.4733 | 1.76 | 20.68 | -1.259 | #N/A | #N/A | 18.131 | -9.717 | #N/A | #N/A | #N/A | -35.1409 | -40.81 | -20.29 | -25.458 | #N/A | 23 | 1 | 0 | 26 |  | 0 |  | 0
24 | Random | 0 | Sample |  |  |  | 3 | 24¦N¦3 | 4 | 24¦C¦4 | 25.453 | 1482 | 25.7 | -6.042 | 122.649 | 5413 | -9.519 |  |  |  |  | -6.042 | -6.042 | 24 | -6.1861 | -5.3234 | -9.519 | -9.5105 | 24 | -9.4851 | -26.6116 | 38.7672 | 275.3708 |  |  | 8.2836 | 7.1032 | [22 empty] | 24.4 | 0.11712 | [8 empty] | 1 | -1.356 | 5 | 6 | -25.366 | 0 | 0 | [12 empty] | 0.25643 | 0.51286 | 0.037883 | 0.075767 | 0 | 24 | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 24 | 24.4 | 25.7 | #N/A |  |  | #N/A | #N/A | 25.453 | -6.0577 | #N/A | #N/A | #N/A | -5.3234 | 1.76 | 20.68 | -1.356 | #N/A | #N/A | 122.649 | -9.717 | #N/A | #N/A | #N/A | -26.6116 | -40.81 | -20.29 | -25.366 | #N/A | 24 | 0 | 1 | 26 | -0.3 | 0 |  | 1
25 | Random | 0 | Sample |  |  |  | 3 | 25¦N¦3 | 4 | 25¦C¦4 | 47.866 | 2784 | 25.6 | 39.073 | 58.45 | 2628 | 15.888 |  |  |  |  | 39.073 | 39.073 | 25 | 38.9436 | 41.0067 | 15.888 | 16.0078 | 25 | 16.0305 | -0.73486 | 75.7914 | 129.1477 |  |  | 1.9872 | 1.704 | [22 empty] | 24.4 | 0.11712 | [8 empty] | 1 | -1.184 | 5 | 6 | -25.384 | 0 | 0 | [12 empty] | 0.2995 | 0.59899 | 0.093129 | 0.18626 | 0 | 25 | 25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 25 | 24.4 | 25.6 | #N/A |  |  | #N/A | #N/A | 47.866 | -6.0577 | #N/A | #N/A | #N/A | 41.0067 | 1.76 | 20.68 | -1.184 | #N/A | #N/A | 58.45 | -9.717 | #N/A | #N/A | #N/A | -0.73486 | -40.81 | -20.29 | -25.384 | #N/A | 25 | 0 | 1 | 26 | -0.4 | 0 |  | 0
26 | Ace Drift5 | 43 | Drift | 0.4 |  | 0.4 | 3 | 26¦N¦3 | 4 | 26¦C¦4 | 27.189 | 1588 | 25.5 | -6.094 | 130.395 | 5756 | -9.431 |  |  | 10.36 | 71.09 | -6.094 | -6.094 | 26 | -6.2086 | -5.3465 | -9.431 | -9.4359 | 26 | -9.4157 | -26.5412 | 41.6349 | 293.0135 | 10.4087 | 73.2534 | 8.2072 | 7.0377 |  |  |  |  | 26 | 24.3 | -6.094 | -9.4359 |  |  |  |  | 41.44 | 27.189 | 284.36 | 130.395 |  |  |  |  | -5.3465 | -26.5412 | 24.3 | 0.10811 | 26 | -6.094 |  |  | -6.094 | -6.094 | -6.094 | -6.2086 | 1 | -1.344 | 5 | 6 | -25.317 | 0 | 0 | 26 | 26 | -0.76873 | 1.5797 | [8 empty] | 0.25657 | 0.51313 | 0.037944 | 0.075889 | 0 | 26 | 26 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 26 | 24.3 | 25.5 | 25.5 | 25.5 | 26 | #N/A | #N/A | #N/A | #N/A | -6.094 | #N/A | -6.2086 | #N/A | 1.76 | 20.68 | -1.344 | #N/A | #N/A | #N/A | #N/A | -9.4359 | #N/A | -9.4157 | #N/A | -40.81 | -20.29 | -25.317 | #N/A | 26 | 0 | 0 | 25.5 |  | 0 |  | 0
27 | Random | 0 | Sample |  |  |  | 3 | 27¦N¦3 | 4 | 27¦C¦4 | 47.171 | 2747 | 25.4 | 0.793 | 57.484 | 2617 | -23.581 |  |  |  |  | 0.793 | 0.793 | 27 | 0.69312 | 1.7388 | -23.581 | -23.4596 | 27 | -23.442 | -40.7659 | 74.6434 | 126.9475 |  |  | 1.9833 | 1.7007 | [22 empty] | 24.2 | 0.099099 | [8 empty] | 1 | -1.175 | 5 | 6 | -25.375 | 0 | 0 | [12 empty] | 0.21982 | 0.43965 | 0.046602 | 0.093204 | 1 | 27 | 27 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 27 | 24.2 | 25.4 | #N/A |  |  | #N/A | #N/A | 47.171 | -6.0577 | #N/A | #N/A | #N/A | 1.7388 | 1.76 | 20.68 | -1.175 | #N/A | #N/A | 57.484 | -9.717 | #N/A | #N/A | #N/A | -40.7659 | -40.81 | -20.29 | -25.375 | #N/A | 27 | 0 | 0 | 25.4 |  | 0 |  | 0 |  | 27
28 | Random | 0 | Sample |  |  |  | 3 | 28¦N¦3 | 4 | 28¦C¦4 | 22.944 | 1325 | 25.3 | -6.26 | 71.395 | 3218 | -9.387 |  |  |  |  | -6.26 | -6.26 | 28 | -6.3451 | -5.4866 | -9.387 | -9.2897 | 28 | -9.2747 | -26.3982 | 34.6226 | 158.632 |  |  | 5.3432 | 4.5818 | [22 empty] | 24.1 | 0.09009 | [8 empty] | 1 | -1.275 | 5 | 6 | -25.399 | 0 | 0 | [12 empty] | 0.25739 | 0.51478 | 0.038072 | 0.076143 | 2 | 28 | 28 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 28 | 24.1 | 25.3 | #N/A |  |  | #N/A | #N/A | 22.944 | -6.0577 | #N/A | #N/A | #N/A | -5.4866 | 1.76 | 20.68 | -1.275 | #N/A | #N/A | 71.395 | -9.717 | #N/A | #N/A | #N/A | -26.3982 | -40.81 | -20.29 | -25.399 | #N/A | 28 | 0 | 0 | 25.3 |  | 0 |  | 0 |  | 27
29 | Random | 0 | Sample |  |  |  | 3 | 29¦N¦3 | 4 | 29¦C¦4 | 50.397 | 3000 | 25.1 | -5.694 | 240.674 | 10125 | -9.519 |  |  |  |  | -5.694 | -5.694 | 29 | -5.7644 | -4.8904 | -9.519 | -9.715 | 29 | -9.7026 | -26.8321 | 79.9724 | 544.1909 |  |  | 7.9356 | 6.8047 | [22 empty] | 24.1 | 0.09009 | [8 empty] | 1 | -1.3 | 5 | 6 | -25.35 | 0 | 0 | [12 empty] | 0.25391 | 0.50782 | 0.037698 | 0.075397 | 3 | 29 | 29 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 29 | 24.1 | 25.1 | #N/A |  |  | #N/A | #N/A | 50.397 | -6.0577 | #N/A | #N/A | #N/A | -4.8904 | 1.76 | 20.68 | -1.3 | #N/A | #N/A | 240.674 | -9.717 | #N/A | #N/A | #N/A | -26.8321 | -40.81 | -20.29 | -25.35 | #N/A | 29 | 1 | 0 | 25.1 |  | 0 |  | 0 |  | 27
30 | Ace Drift6 | 43 | Drift | 0.418 |  | 0.418 | 3 | 30¦N¦3 | 4 | 30¦C¦4 | 27.844 | 1626 | 25.1 | -6.19 | 134.66 | 5934 | -9.478 |  |  | 10.36 | 71.09 | -6.19 | -6.19 | 30 | -6.2456 | -5.3845 | -9.478 | -9.4903 | 30 | -9.4805 | -26.6069 | 42.7169 | 302.7277 | 10.2194 | 72.4229 | 8.2645 | 7.0868 |  |  |  |  | 30 | 24.1 | -6.19 | -9.4903 |  |  |  |  | 43.3048 | 27.844 | 297.1562 | 134.66 |  |  |  |  | -5.3845 | -26.6069 | 24.1 | 0.09009 |  |  | 30 | -6.19 | -6.19 | -6.19 | -6.19 | -6.2456 | 1 | -1.167 | 5 | 6 | -25.352 | 0 | 0 | 30 | 30 | -1.052 | -0.034001 | [8 empty] | 0.25679 | 0.51358 | 0.037887 | 0.075775 | 3 | 30 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 30 | 24.1 | 25.1 | 25.1 | 25.1 | 30 | #N/A | #N/A | #N/A | #N/A | -6.19 | #N/A | -6.2456 | #N/A | 1.76 | 20.68 | -1.167 | #N/A | #N/A | #N/A | #N/A | -9.4903 | #N/A | -9.4805 | #N/A | -40.81 | -20.29 | -25.352 | #N/A | 30 | 0 | 1 | 25.1 | 0 | 0 |  | 0 |  | 27
31 | Random | 0 | Sample |  |  |  | 3 | 31¦N¦3 | 4 | 31¦C¦4 | 25.209 | 1470 | 25 | -5.844 | 121.954 | 5416 | -9.642 |  |  |  |  | -5.844 | -5.844 | 31 | -5.8849 | -5.0141 | -9.642 | -9.6323 | 31 | -9.6251 | -26.7535 | 38.3641 | 273.7879 |  |  | 8.3225 | 7.1366 | [22 empty] | 23.9 | 0.072072 | [8 empty] | 1 | -1.151 | 5 | 6 | -25.396 | 0 | 0 | [12 empty] | 0.25463 | 0.50925 | 0.037763 | 0.075526 | 3 | 31 | 31 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 31 | 23.9 | 25 | #N/A |  |  | #N/A | #N/A | 25.209 | -6.0577 | #N/A | #N/A | #N/A | -5.0141 | 1.76 | 20.68 | -1.151 | #N/A | #N/A | 121.954 | -9.717 | #N/A | #N/A | #N/A | -26.7535 | -40.81 | -20.29 | -25.396 | #N/A | 31 | 0 | 1 | 25.1 | -0.1 | 0 |  | 0 |  | 27
32 | Random | 0 | Sample |  |  |  | 3 | 32¦N¦3 | 4 | 32¦C¦4 | 12.229 | 654 | 24.9 | -6.39 | 56.107 | 2542 | -9.702 |  |  |  |  | -6.39 | -6.39 | 32 | -6.4161 | -5.5595 | -9.702 | -9.5782 | 32 | -9.5736 | -26.7013 | 16.9224 | 123.8112 |  |  | 8.5323 | 7.3164 | [22 empty] | 23.9 | 0.072072 | [8 empty] | 1 | -1.253 | 5 | 6 | -25.392 | 0 | 0 | [12 empty] | 0.25782 | 0.51564 | 0.037807 | 0.075614 | 4 | 32 | 32 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 32 | 23.9 | 24.9 | #N/A |  |  | #N/A | #N/A | 12.229 | -6.0577 | #N/A | #N/A | #N/A | -5.5595 | 1.76 | 20.68 | -1.253 | #N/A | #N/A | 56.107 | -9.717 | #N/A | #N/A | #N/A | -26.7013 | -40.81 | -20.29 | -25.392 | #N/A | 32 | 0 | 0 | 24.9 |  | 0 |  | 0 |  | 27
33 | Random | 0 | Sample |  |  |  | 3 | 33¦N¦3 | 4 | 33¦C¦4 | 64.233 | 3810 | 24.7 | -5.821 | 302.204 | 12392 | -9.333 |  |  |  |  | -5.821 | -5.821 | 33 | -5.8323 | -4.9602 | -9.333 | -9.6356 | 33 | -9.6336 | -26.7622 | 102.8282 | 684.335 |  |  | 7.7611 | 6.6551 | [22 empty] | 23.7 | 0.054054 | [8 empty] | 1 | -1.263 | 5 | 6 | -25.314 | 0 | 0 | [12 empty] | 0.25431 | 0.50862 | 0.037756 | 0.075512 | 5 | 33 | 33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 33 | 23.7 | 24.7 | #N/A |  |  | #N/A | #N/A | 64.233 | -6.0577 | #N/A | #N/A | #N/A | -4.9602 | 1.76 | 20.68 | -1.263 | #N/A | #N/A | 302.204 | -9.717 | #N/A | #N/A | #N/A | -26.7622 | -40.81 | -20.29 | -25.314 | #N/A | 33 | 1 | 0 | 24.7 |  | 0 |  | 0 |  | 27
34 | Ace Drift7 | 43 | Drift | 0.399 |  | 0.399 | 3 | 34¦N¦3 | 4 | 34¦C¦4 | 27.2 | 1544 | 24.8 | -6.213 | 127.619 | 5661 | -9.469 |  |  | 10.36 | 71.09 | -6.213 | -6.213 | 34 | -6.2096 | -5.3475 | -9.469 | -9.4691 | 34 | -9.4697 | -26.5959 | 41.6531 | 286.6908 | 10.4394 | 71.8523 | 8.0266 | 6.8828 |  |  |  |  | 34 | 23.9 | -6.213 | -9.4691 |  |  |  |  | 41.3364 | 27.2 | 283.6491 | 127.619 |  |  |  |  | -5.3475 | -26.5959 | 23.9 | 0.072072 |  |  | 34 | -6.213 | -6.213 | -6.213 | -6.213 | -6.2096 | 1 | -1.205 | 5 | 6 | -25.34 | 0 | 0 | 34 | 34 | -0.76641 | 0.2052 | [8 empty] | 0.25657 | 0.51314 | 0.037897 | 0.075794 | 5 | 34 | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 34 | 23.9 | 24.8 | 24.8 | 24.8 | 34 | #N/A | #N/A | #N/A | #N/A | -6.213 | #N/A | -6.2096 | #N/A | 1.76 | 20.68 | -1.205 | #N/A | #N/A | #N/A | #N/A | -9.4691 | #N/A | -9.4697 | #N/A | -40.81 | -20.29 | -25.34 | #N/A | 34 | 0 | 1 | 24.7 | 0.1 | 0 |  | 0 |  | 27
35 | Random | 0 | Sample |  |  |  | 3 | 35¦N¦3 | 4 | 35¦C¦4 | 23.414 | 1317 | 24.6 | -6.033 | 109.657 | 4857 | -9.637 |  |  |  |  | -6.033 | -6.033 | 35 | -6.0148 | -5.1476 | -9.637 | -9.606 | 35 | -9.6092 | -26.7374 | 35.399 | 245.7796 |  |  | 8.0969 | 6.9431 | [22 empty] | 23.7 | 0.054054 | [8 empty] | 1 | -1.269 | 5 | 6 | -25.345 | 0 | 0 | [12 empty] | 0.2554 | 0.5108 | 0.037777 | 0.075553 | 5 | 35 | 35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 35 | 23.7 | 24.6 | #N/A |  |  | #N/A | #N/A | 23.414 | -6.0577 | #N/A | #N/A | #N/A | -5.1476 | 1.76 | 20.68 | -1.269 | #N/A | #N/A | 109.657 | -9.717 | #N/A | #N/A | #N/A | -26.7374 | -40.81 | -20.29 | -25.345 | #N/A | 35 | 0 | 1 | 24.7 | -0.1 | 0 |  | 0 |  | 27
36 | Random | 0 | Sample |  |  |  | 3 | 36¦N¦3 | 4 | 36¦C¦4 | 22.876 | 1292 | 24.5 | -6.116 | 107.199 | 4715 | -9.529 |  |  |  |  | -6.116 | -6.116 | 36 | -6.0831 | -5.2176 | -9.529 | -9.4937 | 36 | -9.4995 | -26.6262 | 34.5102 | 240.1811 |  |  | 8.1163 | 6.9597 | [22 empty] | 23.6 | 0.045045 | [8 empty] | 1 | -1.155 | 5 | 6 | -25.39 | 0 | 0 | [12 empty] | 0.25581 | 0.51162 | 0.037871 | 0.075742 | 6 | 36 | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 36 | 23.6 | 24.5 | #N/A |  |  | #N/A | #N/A | 22.876 | -6.0577 | #N/A | #N/A | #N/A | -5.2176 | 1.76 | 20.68 | -1.155 | #N/A | #N/A | 107.199 | -9.717 | #N/A | #N/A | #N/A | -26.6262 | -40.81 | -20.29 | -25.39 | #N/A | 36 | 0 | 0 | 24.5 |  | 0 |  | 0 |  | 27
37 | Random | 0 | Sample |  |  |  | 3 | 37¦N¦3 | 4 | 37¦C¦4 | 31.222 | 1790 | 24.5 | -5.902 | 146.952 | 6395 | -9.55 |  |  |  |  | -5.902 | -5.902 | 37 | -5.8543 | -4.9828 | -9.55 | -9.5836 | 37 | -9.592 | -26.72 | 48.2971 | 330.7247 |  |  | 7.9857 | 6.8477 | [22 empty] | 23.5 | 0.036036 | [8 empty] | 1 | -1.171 | 5 | 6 | -25.339 | 0 | 0 | [12 empty] | 0.25444 | 0.50889 | 0.037791 | 0.075582 | 7 | 37 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 37 | 23.5 | 24.5 | #N/A |  |  | #N/A | #N/A | 31.222 | -6.0577 | #N/A | #N/A | #N/A | -4.9828 | 1.76 | 20.68 | -1.171 | #N/A | #N/A | 146.952 | -9.717 | #N/A | #N/A | #N/A | -26.72 | -40.81 | -20.29 | -25.339 | #N/A | 37 | 0 | 0 | 24.5 |  | 0 |  | 0 |  | 27
38 | Ace Drift8 | 43 | Drift | 0.394 |  | 0.394 | 3 | 38¦N¦3 | 4 | 38¦C¦4 | 27.2 | 1545 | 24.5 | -6.426 | 127.597 | 5603 | -9.426 |  |  | 10.36 | 71.09 | -6.426 | -6.426 | 38 | -6.3636 | -5.5056 | -9.426 | -9.4261 | 38 | -9.437 | -26.5628 | 41.6531 | 286.6407 | 10.5719 | 72.7514 | 8.0252 | 6.8816 |  |  |  |  | 38 | 23.6 | -6.426 | -9.4261 |  |  |  |  | 40.8184 | 27.2 | 280.0946 | 127.597 |  |  |  |  | -5.5056 | -26.5628 | 23.6 | 0.045045 |  |  | 38 | -6.426 | -6.426 | -6.426 | -6.426 | -6.3636 | 1 | -1.186 | 5 | 6 | -25.377 | 0 | 0 | 38 | 38 | -2.2302 | 0.97173 | [8 empty] | 0.2575 | 0.515 | 0.037926 | 0.075851 | 8 | 38 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 38 | 23.6 | 24.5 | 24.5 | 24.5 | 38 | #N/A | #N/A | #N/A | #N/A | -6.426 | #N/A | -6.3636 | #N/A | 1.76 | 20.68 | -1.186 | #N/A | #N/A | #N/A | #N/A | -9.4261 | #N/A | -9.437 | #N/A | -40.81 | -20.29 | -25.377 | #N/A | 38 | 0 | 0 | 24.5 |  | 0 |  | 0 |  | 27
39 | Random | 0 | Sample |  |  |  | 3 | 39¦N¦3 | 4 | 39¦C¦4 | 13.845 | 747 | 24.4 | -6.243 | 62.657 | 2836 | -9.567 |  |  |  |  | -6.243 | -6.243 | 39 | -6.1658 | -5.3026 | -9.567 | -9.4545 | 39 | -9.4681 | -26.5943 | 19.5919 | 138.7298 |  |  | 8.2577 | 7.081 | [22 empty] | 23.5 | 0.036036 | [8 empty] | 1 | -1.196 | 5 | 6 | -25.386 | 0 | 0 | [12 empty] | 0.25631 | 0.51262 | 0.037898 | 0.075797 | 9 | 39 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 39 | 23.5 | 24.4 | #N/A |  |  | #N/A | #N/A | 13.845 | -6.0577 | #N/A | #N/A | #N/A | -5.3026 | 1.76 | 20.68 | -1.196 | #N/A | #N/A | 62.657 | -9.717 | #N/A | #N/A | #N/A | -26.5943 | -40.81 | -20.29 | -25.386 | #N/A | 39 | 0 | 0 | 24.4 |  | 0 |  | 0 |  | 27
40 | Random | 0 | Sample |  |  |  | 3 | 40¦N¦3 | 4 | 40¦C¦4 | 2.809 | 157 | 24.3 | 1.348 | 22.965 | 1041 | -10.776 |  |  |  |  | 1.348 | 1.348 | 40 | 1.4399 | 2.5055 | -10.776 | -10.5948 | 40 | -10.6109 | -27.7533 | 1.3614 | 48.3252 |  |  | 41.3957 | 35.4968 | [22 empty] | 23.4 | 0.027027 | [8 empty] | 1 | -1.158 | 5 | 6 | -25.44 | 0 | 0 | [12 empty] | 0.21653 | 0.43306 | 0.037033 | 0.074066 | 10 | 40 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 40 | 23.4 | 24.3 | #N/A |  |  | #N/A | #N/A | 2.809 | -6.0577 | #N/A | #N/A | #N/A | 2.5055 | 1.76 | 20.68 | -1.158 | #N/A | #N/A | 22.965 | -9.717 | #N/A | #N/A | #N/A | -27.7533 | -40.81 | -20.29 | -25.44 | #N/A | 40 | 0 | 0 | 24.3 |  | 0 |  | 0 |  | 27
41 | Random | 0 | Sample |  |  |  | 3 | 41¦N¦3 | 4 | 41¦C¦4 | 4.359 | 194 | 24.1 | 0.297 | 27.265 | 1238 | -10.813 |  |  |  |  | 0.297 | 0.297 | 41 | 0.40369 | 1.4417 | -10.813 | -10.6392 | 41 | -10.658 | -27.801 | 3.9219 | 58.1191 |  |  | 17.282 | 14.8193 | [22 empty] | 23.2 | 0.009009 | [8 empty] | 1 | -1.195 | 5 | 6 | -25.415 | 0 | 0 | [12 empty] | 0.22114 | 0.44228 | 0.037003 | 0.074007 | 11 | 41 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 41 | 23.2 | 24.1 | #N/A |  |  | #N/A | #N/A | 4.359 | -6.0577 | #N/A | #N/A | #N/A | 1.4417 | 1.76 | 20.68 | -1.195 | #N/A | #N/A | 27.265 | -9.717 | #N/A | #N/A | #N/A | -27.801 | -40.81 | -20.29 | -25.415 | #N/A | 41 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
42 | Ace Drift9 | 43 | Drift | 0.38 |  | 0.38 | 3 | 42¦N¦3 | 4 | 42¦C¦4 | 25.931 | 1469 | 24.1 | -6.24 | 121.336 | 5413 | -9.51 |  |  | 10.36 | 71.09 | -6.24 | -6.24 | 42 | -6.1186 | -5.254 | -9.51 | -9.4992 | 42 | -9.5206 | -26.6476 | 39.5568 | 272.3803 | 10.4097 | 71.679 | 8.0301 | 6.8858 |  |  |  |  | 42 | 23.2 | -6.24 | -9.4992 |  |  |  |  | 39.368 | 25.931 | 270.142 | 121.336 |  |  |  |  | -5.254 | -26.6476 | 23.2 | 0.009009 |  |  | 42 | -6.24 | -6.24 | -6.24 | -6.24 | -6.1186 | 1 | -1.09 | 5 | 6 | -25.347 | 0 | 0 | 42 | 42 | -0.11305 | -0.97459 | [8 empty] | 0.25602 | 0.51205 | 0.037853 | 0.075705 | 12 | 42 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 42 | 23.2 | 24.1 | 24.1 | 24.1 | 42 | #N/A | #N/A | #N/A | #N/A | -6.24 | #N/A | -6.1186 | #N/A | 1.76 | 20.68 | -1.09 | #N/A | #N/A | #N/A | #N/A | -9.4992 | #N/A | -9.5206 | #N/A | -40.81 | -20.29 | -25.347 | #N/A | 42 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
43 | Random | 0 | Sample |  |  |  | 3 | 43¦N¦3 | 4 | 43¦C¦4 | 2.919 | 105 | 24.1 | -1.599 | 15.34 | 710 | -10.334 |  |  |  |  | -1.599 | -1.599 | 43 | -1.4628 | -0.47444 | -10.334 | -10.1395 | 43 | -10.1635 | -27.2996 | 1.5431 | 30.9581 |  |  | 23.3962 | 20.0622 | [22 empty] | 23.3 | 0.018018 | [8 empty] | 1 | -1.277 | 5 | 6 | -25.447 | 0 | 0 | [12 empty] | 0.23015 | 0.4603 | 0.037339 | 0.074677 | 13 | 43 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 23.3 | 24.1 | #N/A |  |  | #N/A | #N/A | 2.919 | -6.0577 | #N/A | #N/A | #N/A | -0.47444 | 1.76 | 20.68 | -1.277 | #N/A | #N/A | 15.34 | -9.717 | #N/A | #N/A | #N/A | -27.2996 | -40.81 | -20.29 | -25.447 | #N/A | 43 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
44 | Random | 0 | Sample |  |  |  | 3 | 44¦N¦3 | 4 | 44¦C¦4 | 2.479 | 143 | 24 | 1.903 | 20.814 | 954 | -10.725 |  |  |  |  | 1.903 | 1.903 | 44 | 2.0539 | 3.1358 | -10.725 | -10.54 | 44 | -10.5666 | -27.7084 | 0.81626 | 43.426 |  |  | 62.0418 | 53.2008 | [22 empty] | 23.2 | 0.009009 | [8 empty] | 1 | -1.265 | 5 | 6 | -25.415 | 0 | 0 | [12 empty] | 0.21394 | 0.42789 | 0.037061 | 0.074123 | 14 | 44 | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 44 | 23.2 | 24 | #N/A |  |  | #N/A | #N/A | 2.479 | -6.0577 | #N/A | #N/A | #N/A | 3.1358 | 1.76 | 20.68 | -1.265 | #N/A | #N/A | 20.814 | -9.717 | #N/A | #N/A | #N/A | -27.7084 | -40.81 | -20.29 | -25.415 | #N/A | 44 | 0 | 0 | 24 |  | 0 |  | 0 |  | 27
45 | Random | 0 | Sample |  |  |  | 3 | 45¦N¦3 | 4 | 45¦C¦4 | 2.478 | 141 | 24 | 1.975 | 20.462 | 945 | -10.909 |  |  |  |  | 1.975 | 1.975 | 45 | 2.1407 | 3.2249 | -10.909 | -10.7234 | 45 | -10.7526 | -27.897 | 0.81461 | 42.6242 |  |  | 61.0199 | 52.3245 | [22 empty] | 23.1 | 0 | [8 empty] | 1 | -1.315 | 5 | 6 | -25.444 | 0 | 0 | [12 empty] | 0.21359 | 0.42717 | 0.036945 | 0.073891 | 15 | 45 | 45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 45 | 23.1 | 24 | #N/A |  |  | #N/A | #N/A | 2.478 | -6.0577 | #N/A | #N/A | #N/A | 3.2249 | 1.76 | 20.68 | -1.315 | #N/A | #N/A | 20.462 | -9.717 | #N/A | #N/A | #N/A | -27.897 | -40.81 | -20.29 | -25.444 | #N/A | 45 | 0 | 0 | 24 |  | 0 |  | 0 |  | 27
46 | Ace Drift10 | 43 | Drift | 0.419 |  | 0.419 | 3 | 46¦N¦3 | 4 | 46¦C¦4 | 29.047 | 1652 | 23.9 | -6.231 | 135.61 | 5971 | -9.407 |  |  | 10.36 | 71.09 | -6.231 | -6.231 | 46 | -6.0505 | -5.1842 | -9.407 | -9.4209 | 46 | -9.4527 | -26.5787 | 44.7042 | 304.8915 | 10.6693 | 72.7665 | 7.9536 | 6.8202 |  |  |  |  | 46 | 23.1 | -6.231 | -9.4209 |  |  |  |  | 43.4084 | 29.047 | 297.8671 | 135.61 |  |  |  |  | -5.1842 | -26.5787 | 23.1 | 0 |  |  | 46 | -6.231 | -6.231 | -6.231 | -6.231 | -6.0505 | 1 | -1.263 | 5 | 6 | -25.35 | 0 | 0 | 46 | 46 | 0.37542 | 0.60837 | [8 empty] | 0.25562 | 0.51123 | 0.037912 | 0.075824 | 16 | 46 | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 46 | 23.1 | 23.9 | 23.9 | 23.9 | 46 | #N/A | #N/A | #N/A | #N/A | -6.231 | #N/A | -6.0505 | #N/A | 1.76 | 20.68 | -1.263 | #N/A | #N/A | #N/A | #N/A | -9.4209 | #N/A | -9.4527 | #N/A | -40.81 | -20.29 | -25.35 | #N/A | 46 | 0 | 0 | 23.9 |  | 0 |  | 0 |  | 27
47 | Random | 0 | Sample |  |  |  | 3 | 47¦N¦3 | 4 | 47¦C¦4 | 2.464 | 142 | 24.1 | 1.617 | 20.21 | 926 | -10.686 |  |  |  |  | 1.617 | 1.617 | 47 | 1.8122 | 2.8877 | -10.686 | -10.5 | 47 | -10.5344 | -27.6757 | 0.79149 | 42.0503 |  |  | 61.9571 | 53.1282 | [22 empty] | 23.3 | 0.018018 | [8 empty] | 1 | -1.175 | 5 | 6 | -25.449 | 0 | 0 | [12 empty] | 0.21495 | 0.4299 | 0.037082 | 0.074164 | 17 | 47 | 47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 47 | 23.3 | 24.1 | #N/A |  |  | #N/A | #N/A | 2.464 | -6.0577 | #N/A | #N/A | #N/A | 2.8877 | 1.76 | 20.68 | -1.175 | #N/A | #N/A | 20.21 | -9.717 | #N/A | #N/A | #N/A | -27.6757 | -40.81 | -20.29 | -25.449 | #N/A | 47 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
48 | Random | 0 | Sample |  |  |  | 3 | 48¦N¦3 | 4 | 48¦C¦4 | 27.231 | 1538 | 23.9 | -6.25 | 127.336 | 5640 | -9.587 |  |  |  |  | -6.25 | -6.25 | 48 | -6.04 | -5.1734 | -9.587 | -9.5866 | 48 | -9.6236 | -26.752 | 41.7043 | 286.0462 |  |  | 7.9987 | 6.8589 | [22 empty] | 23.1 | 0 | [8 empty] | 1 | -1.244 | 5 | 6 | -25.357 | 0 | 0 | [12 empty] | 0.25555 | 0.51111 | 0.037764 | 0.075529 | 18 | 48 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 48 | 23.1 | 23.9 | #N/A |  |  | #N/A | #N/A | 27.231 | -6.0577 | #N/A | #N/A | #N/A | -5.1734 | 1.76 | 20.68 | -1.244 | #N/A | #N/A | 127.336 | -9.717 | #N/A | #N/A | #N/A | -26.752 | -40.81 | -20.29 | -25.357 | #N/A | 48 | 0 | 0 | 23.9 |  | 0 |  | 0 |  | 27
49 | Random | 0 | Sample |  |  |  | 3 | 49¦N¦3 | 4 | 49¦C¦4 | 41.648 | 2404 | 24.1 | -5.971 | 195.21 | 8403 | -9.471 |  |  |  |  | -5.971 | -5.971 | 49 | -5.7463 | -4.8718 | -9.471 | -9.5882 | 49 | -9.6278 | -26.7563 | 65.5199 | 440.6397 |  |  | 7.8429 | 6.7253 | [22 empty] | 23.2 | 0.009009 | [8 empty] | 1 | -1.223 | 5 | 6 | -25.338 | 0 | 0 | [12 empty] | 0.2538 | 0.5076 | 0.037761 | 0.075522 | 19 | 49 | 49 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 49 | 23.2 | 24.1 | #N/A |  |  | #N/A | #N/A | 41.648 | -6.0577 | #N/A | #N/A | #N/A | -4.8718 | 1.76 | 20.68 | -1.223 | #N/A | #N/A | 195.21 | -9.717 | #N/A | #N/A | #N/A | -26.7563 | -40.81 | -20.29 | -25.338 | #N/A | 49 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
50 | Ace Drift11 | 43 | Drift | 0.424 |  | 0.424 | 3 | 50¦N¦3 | 4 | 50¦C¦4 | 28.692 | 1623 | 24.2 | -6.161 | 133.29 | 5892 | -9.488 |  |  | 10.36 | 71.09 | -6.161 | -6.161 | 50 | -5.9215 | -5.0518 | -9.488 | -9.4979 | 50 | -9.5401 | -26.6673 | 44.1177 | 299.6074 | 10.4051 | 70.6621 | 7.9196 | 6.7911 |  |  |  |  | 50 | 23.3 | -6.161 | -9.4979 |  |  |  |  | 43.9264 | 28.692 | 301.4216 | 133.29 |  |  |  |  | -5.0518 | -26.6673 | 23.3 | 0.018018 |  |  | 50 | -6.161 | -6.161 | -6.161 | -6.161 | -5.9215 | 1 | -1.258 | 5 | 6 | -25.373 | 0 | 0 | 50 | 50 | 1.4654 | -1.5957 | [8 empty] | 0.25484 | 0.50969 | 0.037836 | 0.075671 | 20 | 50 | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 50 | 23.3 | 24.2 | 24.2 | 24.2 | 50 | #N/A | #N/A | #N/A | #N/A | -6.161 | #N/A | -5.9215 | #N/A | 1.76 | 20.68 | -1.258 | #N/A | #N/A | #N/A | #N/A | -9.4979 | #N/A | -9.5401 | #N/A | -40.81 | -20.29 | -25.373 | #N/A | 50 | 0 | 0 | 24.2 |  | 0 |  | 0 |  | 27
51 | Random | 0 | Sample |  |  |  | 3 | 51¦N¦3 | 4 | 51¦C¦4 | 16.082 | 855 | 24.1 | -6.109 | 71.442 | 3202 | -9.619 |  |  |  |  | -6.109 | -6.109 | 51 | -5.8548 | -4.9832 | -9.619 | -9.5218 | 51 | -9.5665 | -26.6941 | 23.2872 | 158.739 |  |  | 7.9494 | 6.8166 | [22 empty] | 23.3 | 0.018018 | [8 empty] | 1 | -1.183 | 5 | 6 | -25.407 | 0 | 0 | [12 empty] | 0.25445 | 0.50889 | 0.037813 | 0.075626 | 21 | 51 | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 51 | 23.3 | 24.1 | #N/A |  |  | #N/A | #N/A | 16.082 | -6.0577 | #N/A | #N/A | #N/A | -4.9832 | 1.76 | 20.68 | -1.183 | #N/A | #N/A | 71.442 | -9.717 | #N/A | #N/A | #N/A | -26.6941 | -40.81 | -20.29 | -25.407 | #N/A | 51 | 0 | 0 | 24.1 |  | 0 |  | 0 |  | 27
52 | Random | 0 | Sample |  |  |  | 3 | 52¦N¦3 | 4 | 52¦C¦4 | 22.447 | 1236 | 24 | -5.883 | 101.638 | 4540 | -9.614 |  |  |  |  | -5.883 | -5.883 | 52 | -5.614 | -4.7361 | -9.614 | -9.5691 | 52 | -9.6164 | -26.7448 | 33.8016 | 227.5151 |  |  | 7.8494 | 6.7309 | [22 empty] | 23.2 | 0.009009 | [8 empty] | 1 | -1.193 | 5 | 6 | -25.377 | 0 | 0 | [12 empty] | 0.25302 | 0.50603 | 0.03777 | 0.075541 | 22 | 52 | 52 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 52 | 23.2 | 24 | #N/A |  |  | #N/A | #N/A | 22.447 | -6.0577 | #N/A | #N/A | #N/A | -4.7361 | 1.76 | 20.68 | -1.193 | #N/A | #N/A | 101.638 | -9.717 | #N/A | #N/A | #N/A | -26.7448 | -40.81 | -20.29 | -25.377 | #N/A | 52 | 0 | 0 | 24 |  | 0 |  | 0 |  | 27
53 | Random | 0 | Sample |  |  |  | 3 | 53¦N¦3 | 4 | 53¦C¦4 | 1.4 | 63 | 25.7 | -9.347 | [7 empty] | -9.347 | -9.347 | 53 | -9.0633 | -8.2771 | [5 empty] | -0.96614 | [27 empty] | 23.2 | 0.009009 | [8 empty] | 1 | -1.245 | 4 | 5 | -25.479 | 0 | 0 | [12 empty] | 0.2744 | 0.54881 |  |  | 23 | 53 | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 53 | 23.2 | 25.7 | #N/A |  |  | #N/A | #N/A | 1.4 | -6.0577 | #N/A | #N/A | #N/A | -8.2771 | 1.76 | 20.68 | -1.245 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.479 | #N/A | 53 | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
54 | Ace Drift12 | 43 | Drift | 0.38 |  | 0.38 | 3 | 54¦N¦3 | 4 | 54¦C¦4 | 27.037 | 1494 | 23.8 | -6.297 | 122.82 | 5383 | -9.406 |  |  | 10.36 | 71.09 | -6.297 | -6.297 | 54 | -5.9985 | -5.1308 | -9.406 | -9.3978 | 54 | -9.4503 | -26.5763 | 41.3838 | 275.7603 | 10.8905 | 72.5685 | 7.7708 | 6.6635 |  |  |  |  | 54 | 23 | -6.297 | -9.3978 |  |  |  |  | 39.368 | 27.037 | 270.142 | 122.82 |  |  |  |  | -5.1308 | -26.5763 | 23 | -0.009009 |  |  | 54 | -6.297 | -6.297 | -6.297 | -6.297 | -5.9985 | 1 | -1.255 | 5 | 6 | -25.342 | 0 | 0 | 54 | 54 | 0.8154 | 0.70677 | [8 empty] | 0.2553 | 0.51061 | 0.037914 | 0.075828 | 24 | 54 | 54 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 54 | 23 | 23.8 | 23.8 | 23.8 | 54 | #N/A | #N/A | #N/A | #N/A | -6.297 | #N/A | -5.9985 | #N/A | 1.76 | 20.68 | -1.255 | #N/A | #N/A | #N/A | #N/A | -9.3978 | #N/A | -9.4503 | #N/A | -40.81 | -20.29 | -25.342 | #N/A | 54 | 0 | 0 | 23.8 |  | 0 |  | 0 |  | 27
55 | Ace Drift 13 | 43 | Drift | 0.397 |  | 0.397 | 3 | 55¦N¦3 | 4 | 55¦C¦4 | 28.314 | 1555 | 24 | -6.395 | 128.087 | 5635 | -9.436 |  |  | 10.36 | 71.09 | -6.395 | -6.395 | 55 | -6.0818 | -5.2162 | -9.436 | -9.4369 | 55 | -9.4921 | -26.6186 | 43.4933 | 287.7567 | 10.9555 | 72.4828 | 7.7156 | 6.6161 |  |  |  |  | 55 | 23.2 | -6.395 | -9.4369 |  |  |  |  | 41.1292 | 28.314 | 282.2273 | 128.087 |  |  |  |  | -5.2162 | -26.6186 | 23.2 | 0.009009 |  |  | 55 | -6.395 | -6.395 | -6.395 | -6.395 | -6.0818 | 1 | -1.346 | 5 | 6 | -25.343 | 0 | 0 | 55 | 55 | 0.1595 | -0.31995 | [8 empty] | 0.2558 | 0.51161 | 0.037877 | 0.075755 | 25 | 55 | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 55 | 23.2 | 24 | 24 | 24 | 55 | #N/A | #N/A | #N/A | #N/A | -6.395 | #N/A | -6.0818 | #N/A | 1.76 | 20.68 | -1.346 | #N/A | #N/A | #N/A | #N/A | -9.4369 | #N/A | -9.4921 | #N/A | -40.81 | -20.29 | -25.343 | #N/A | 55 | 0 | 0 | 24 |  | 0 |  | 0 |  | 27
56 | blank capsule | 45 | Blank | 0.397 |  | 0.397 | 3 | 56¦N¦3 | 4 | 56¦C¦4 | 2.058 | 92 | 25.7 | -6.777 | [7 empty] | -6.777 | -6.777 | 56 | -6.449 | -5.5933 | [5 empty] | 0.12081 |  | 0.030431 | [19 empty] | -6.777 | 2.058 |  |  |  |  | 23.2 | 0.009009 | [8 empty] | 1 | -1.246 | 4 | 5 | -25.445 | 0 | 0 | [12 empty] | 0.25802 | 0.51603 |  |  | 26 | 56 | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 45 | 0 | 0 | 0 | 0 | 56 | 23.2 | 25.7 | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | 1.76 | 20.68 | -1.246 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -40.81 | -20.29 | -25.445 | #N/A | 56 | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
57 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
58 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
59 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
60 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
61 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
62 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
63 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
64 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
65 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
66 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
67 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
68 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
69 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
70 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
71 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
72 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
73 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
74 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
75 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
76 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
77 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
78 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
79 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
80 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
81 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
82 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
83 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
84 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
85 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
86 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
87 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
88 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
89 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
90 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
91 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
92 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
93 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
94 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
95 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
96 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
97 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
98 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
99 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
100 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
101 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
102 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
103 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
104 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
105 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
106 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
107 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
108 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
109 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
110 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
111 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
112 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
113 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
114 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
115 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
116 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
117 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
118 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
119 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27
120 | [92 empty] | 26 | [15 empty] | #N/A | #N/A | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A |  | 0 | 0 | 25.7 |  | 0 |  | 0 |  | 27

## Fits
Fits
Every coefficient below is fitted from this run. Nothing here is typed in by hand.
Blank, from the blank capsules
n with a nitrogen peak | 2 | No | Nitrogen blank correction in force? On applies it whenever a blank exists; Auto also requires three blanks and an SD of 2 permil or better.
mean Area 28 | 1.867
mean d15N (permil) | 4.502
SD of blank d15N (permil) | 15.9509 |  | A large SD here means the blank correction is not supported.
n with a carbon peak | 0 | No | Carbon blank correction in force?
mean Area 44 | 0
mean d13C (permil) | 0
SD of blank d13C (permil)
Linearity, nitrogen: d15N against Area 28
n linearity standards | 4
pivot, mean Area 28 | 27.5413 |  | The pivot cancels in the calibration and does not affect the result.
fitted slope (permil per area unit) | -0.056964
r-squared | 0.87908
residual SD (permil) | 0.52896
slope applied | 0 |  | Zero when the correction is off, not supported by the data on Auto, or fewer than three standards were found. Row 92 onwards has the test.
Linearity, carbon: d13C against Area 44
n linearity standards | 4
pivot, mean Area 44 | 127.5635
fitted slope | 0.0017328
r-squared | 0.96075
residual SD (permil) | 0.041939
slope applied | 0.0017328
Drift, nitrogen: linearity-corrected d15N of the drift standards
n drift standards | 13
slope against row | -0.014754
r-squared against row | 0.72977
residual SD against row | 0.14264
slope against background | 0.15644
r-squared against background | 0.87348
residual SD against background | 0.097601
model used | Row |  | On Auto, whichever model left the smaller residual SD. That is not an independent test, because both models were fitted to these same standards. Judge by the check standards on QC.
slope applied | -0.014754
pivot applied | 33.7692 |  | Mean of the x values used in the fit.
Drift, carbon: linearity-corrected d13C of the drift standards
n drift standards | 13
slope against row | 0.0025977
r-squared against row | 0.45743
residual SD against row | 0.044946
slope against background | -0.026323 |  | The mass-28 background stands in for a mass-44 background the export does not carry.
r-squared against background | 0.50006
residual SD against background | 0.043144
model used | Row
slope applied | 0.0025977
pivot applied | 33.7692
Calibration, nitrogen: accepted regressed on measured
n anchor measurements | 6
slope | 1.0266
intercept | 1.0273
r-squared | 0.99978
residual SD (permil) | 0.17333 |  | Applied as accepted = slope x measured + intercept. Never inverted.
Calibration, carbon: accepted regressed on measured
n anchor measurements | 6
slope | 1.0142
intercept | -16.9922
r-squared | 1
residual SD (permil) | 0.02084
Elemental k-factor, nitrogen: Area 28 against micrograms N
n standards used | 23
slope, area per ug N | 0.60536
intercept, area blank | 1.9849 |  | Subtracted from every peak, so the k-factor is not forced through the origin.
r-squared | 0.99702
Elemental k-factor, carbon: Area 44 against micrograms C
n standards used | 23
slope, area per ug C | 0.43905
intercept, area blank | 1.7479
r-squared | 0.99683
Size limits used by the flags
minimum trusted Area 28 | 8.889 |  | Defaults to the smallest linearity standard in the run.
minimum trusted Area 44 | 37.895
background floor | 23.1 |  | Mean of the five lowest reference-peak backgrounds in the run.
Instrument settling, from the mass-28 background on the reference peaks
starting excess above the floor | 11.1 |  | Highest of the first five reference-peak backgrounds, less the floor. Under 1 mV the run counts as settled from the start.
settled from position | 27 |  | First of three readings in a row at or under Setup B26, skipping readings within two positions of an oversized sample. A later spike does not move it. One past the last position means the background never settled.
drift slope before that position | -0.0318
n before | 5
drift slope from that position | -0.0028939
n from | 8 |  | A run that settled shows a steep slope before and a flat one after.
Linearity, nitrogen: is the slope real?
SE of slope | 0.014939
t | -3.8131
p, two-sided | 0.062407 |  | Probability of a slope this large from a flat response. Auto applies the slope only below the significance level in Setup.
in force | No
maximum trusted Area 28 | 51.069 |  | Defaults to the largest linearity standard in the run.
chart rug level | -6.0577 |  | Where the QC linearity chart draws the unknowns' sizes, just under the standards.
Linearity, carbon: is the slope real?
SE of slope | 0.00024766
t | 6.9965
p, two-sided | 0.019823
in force | Yes
maximum trusted Area 44 | 240.154
chart rug level | -9.717
Drift, nitrogen: is the slope real?
n monitors used | 13 |  | Drift standards from the first position in Setup B14, minus any listed in Setup B16.
mean x used | 33.7692
Sxx | 2776.3077
intercept, model in use | -5.6043
residual SD, model in use | 0.14264
SE of slope | 0.0027071
t | -5.4503
p, two-sided | 0.00020076
in force | Yes
residual sum of squares | 0.2238
monitors flagged as outliers | 0 |  | Each monitor against the line through the others, in units of their scatter (Work column CB).
Drift, carbon: is the slope real?
n monitors used | 13
mean x used | 33.7692
Sxx | 2776.3077
intercept, model in use | -9.5667
residual SD, model in use | 0.044946
SE of slope | 0.00085302
t | 3.0453
p, two-sided | 0.011141
in force | Yes
residual sum of squares | 0.022222
monitors flagged as outliers | 0
Calibration, nitrogen: statistics behind the uncertainty
mean measured, all anchors | 9.9287
Sxx | 509.3727
pooled SD of anchor replicates | 0.16882 |  | Within-material scatter of the anchors, pooled. Stands in for the repeatability of a single sample analysis when Setup gives none.
sample repeatability in use | 0.16882
its source | pooled anchor replicates
accepted d15N, lowest anchor | 1.76
accepted d15N, highest anchor | 20.68 |  | A sample outside this range is extrapolated beyond the anchors and is flagged.
Calibration, carbon: statistics behind the uncertainty
mean measured, all anchors | -13.3686
Sxx | 614.1011
pooled SD of anchor replicates | 0.020549
sample repeatability in use | 0.020549
its source | pooled anchor replicates
accepted d13C, lowest anchor | -40.81
accepted d13C, highest anchor | -20.29
Uncertainty table, nitrogen (rows 163 to 175)
Material | n | sum x | sum x^2 | mean measured | SD | u(mean) | S1 | S2 | u(accepted)
USGS 64 | 3 | 2.1474 | 1.6205 | 0.71581 | 0.2042 | 0.1179 | -27.6385 | 28.3673 | 0.06
USGS 65 | 3 | 57.4245 | 1099.2209 | 19.1415 | 0.12369 | 0.071414 | 27.6385 | -28.3673 | 0.06
USGS 66 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0.06
USGS 40 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
USGS 41a | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Lin | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Drift | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Bypass | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Blank Capsule | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Blank Hole | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
One row per line of the standards table. n, mean and SD are of the anchor replicates after drift correction; S1 and S2 are the sums the uncertainty weights need; u(accepted) is the 1-sigma certificate value from Setup.
Uncertainty table, carbon (rows 183 to 195)
Material | n | sum x | sum x^2 | mean measured | SD | u(mean) | S1 | S2 | u(accepted)
USGS 64 | 3 | -70.4562 | 1654.6927 | -23.4854 | 0.020561 | 0.011871 | -30.3504 | 30.7798 | 0.04
USGS 65 | 3 | -9.7553 | 31.7229 | -3.2518 | 0.020536 | 0.011857 | 30.3504 | -30.7798 | 0.04
USGS 66 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0.04
USGS 40 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
USGS 41a | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Lin | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Drift | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Ace Bypass | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Blank Capsule | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
Blank Hole | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
| 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0
One row per line of the standards table. n, mean and SD are of the anchor replicates after drift correction; S1 and S2 are the sums the uncertainty weights need; u(accepted) is the 1-sigma certificate value from Setup.
Reference pulses: did the working-gas measurement hold across the run?
The designated reference pulse carries the assigned value by definition, so only the other pulses can show movement. Their d values are relative to the designated pulse of the same acquisition.
N2 pulses read | 56
mean d, relative to the designated pulse | -1.2191
SD (permil) | 0.066341
range (permil) | 0.285
trend, permil per position | -0.00084699
p of the trend, two-sided | 0.12356
CO2 pulses read | 56
mean d, relative to the designated pulse | -25.3937
SD (permil) | 0.065478
range (permil) | 0.422
trend, permil per position | 0.00088161
p of the trend, two-sided | 0.10392
Check standards
check standards with a calibrated value | 0 |  | Materials with the Check role that were measured in this run. Zero means nothing independent of the anchors tests the calibration.
Sample-peak background on the drift monitors
The mass-28 background under each monitor's own peak, from column BGD 28. Present in every export; the reference-peak background above needs the reference peaks. A monitor run right after a large sample can read high from what that sample left behind.
monitors with a background | 13
lowest | 23.8
highest | 30.3
trend, per position | -0.11963
p of the trend, two-sided | 2.9894e-05
Carryover after oversized samples
rise that flags a position (mV) | 2 |  | Setup B28. A position is exposed when one of the two before it is above the linearity range (B97, B104); its rise is its sample-peak background less the last one that was not exposed.
positions above the linearity range | 4
of those, followed directly by a sample or check | 1
positions flagged for possible carryover | 0
largest rise (mV) | 0.1

## Results
Results
Read QC before using these. Isotope values are on the scale set by the anchor standards in Setup; %C, %N and C:N need a mass in Work column F or G. u is the standard uncertainty from the scale calibration and the sample repeatability (Fits rows 132 onwards), and U is k times u with k from Setup B32; neither includes the blank, linearity or drift corrections or day-to-day reproducibility.
Row | Identifier | Type | Mass (mg) | %C | %N | C:N molar | C:N mass | d13C (permil) | d15N (permil) | Flags | u(d13C) | U(d13C), k=2.0 | u(d15N) | U(d15N), k=2.0
1 | Blank Hole | Exclude | [6 empty] | 4.3581 | no C peak; N below linearity range; ran before the background settled; |  |  | 0.20925 | 0.41849
2 | Ace Bypass | Exclude | 1.131 | 72.6795 | 11.3853 | 7.4444 | 6.3836 | -26.5575 | -4.5756 | N above linearity range; C above linearity range; ran before the background settled; | 0.03793 | 0.07586 | 0.25209 | 0.50419
3 | Ace Bypass | Exclude | 0.49 | 71.6518 | 11.0726 | 7.5465 | 6.4711 | -26.5447 | -4.0739 | ran before the background settled; | 0.037941 | 0.075883 | 0.24924 | 0.49848
4 | Blank Capsule | Blank | [6 empty] | 16.7771 | no C peak; N below linearity range; ran before the background settled; |  |  | 0.19096 | 0.38192
5 | Blank Capsule | Blank | [7 empty] | no N peak; no C peak; ran before the background settled;
6 | Ace Lin1 | Linearity | 0.126 | 65.3417 | 9.0516 | 8.4184 | 7.2188 | -26.5643 | -2.4978 | ran before the background settled; | 0.037924 | 0.075849 | 0.24056 | 0.48113
7 | Ace Lin2 | Linearity | 0.177 | 67.3286 | 9.4011 | 8.352 | 7.1618 | -26.5964 | -3.5893 | ran before the background settled; | 0.037897 | 0.075793 | 0.24652 | 0.49304
8 | Ace Lin3 | Linearity | 0.572 | 70.2358 | 10.444 | 7.8425 | 6.725 | -26.53 | -4.9878 | ran before the background settled; | 0.037954 | 0.075909 | 0.25447 | 0.50894
9 | Ace Lin4 | Linearity | 0.788 | 68.9095 | 10.2897 | 7.8099 | 6.697 | -26.6092 | -5.1605 | ran before the background settled; | 0.037885 | 0.075771 | 0.25548 | 0.51095
10 | Ace Drift1 | Drift | 0.4 | 70.2742 | 10.0829 | 8.1279 | 6.9696 | -26.6126 | -5.0027 | ran before the background settled; | 0.037883 | 0.075765 | 0.25456 | 0.50912
11 | USGS 64 | Anchor | 0.397 | 31.2498 | 18.6941 | 1.9494 | 1.6716 | -40.7946 | 1.6587 | ran before the background settled; | 0.046654 | 0.093308 | 0.22018 | 0.44035
12 | USGS 64 | Anchor | 0.409 | 31.2891 | 18.7821 | 1.9427 | 1.6659 | -40.8337 | 2.0034 | ran before the background settled; | 0.046725 | 0.093451 | 0.21867 | 0.43734
13 | USGS 64 | Anchor | 0.389 | 31.4686 | 18.7647 | 1.9557 | 1.677 | -40.8016 | 1.6243 | ran before the background settled; | 0.046667 | 0.093334 | 0.22033 | 0.44066
14 | Ace Drift2 | Drift | 0.428 | 71.9271 | 10.3114 | 8.1347 | 6.9755 | -26.6752 | -5.1689 | ran before the background settled; | 0.037829 | 0.075658 | 0.25553 | 0.51105
15 | USGS 65 | Anchor | 0.402 | 31.4022 | 18.6588 | 1.9626 | 1.683 | -20.2662 | 20.7668 | ran before the background settled; | 0.046722 | 0.093444 | 0.19772 | 0.39544
16 | USGS 65 | Anchor | 0.43 | 31.4677 | 18.506 | 1.983 | 1.7004 | -20.2992 | 20.5324 | ran before the background settled; | 0.046662 | 0.093324 | 0.19716 | 0.39432
17 | USGS 65 | Anchor | 0.423 | 30.7145 | 18.1835 | 1.9698 | 1.6891 | -20.3047 | 20.7344 | ran before the background settled; | 0.046652 | 0.093303 | 0.19764 | 0.39528
18 | Ace Drift3 | Drift | 0.415 | 72.0271 | 10.218 | 8.2205 | 7.049 | -26.547 | -5.3147 | ran before the background settled; | 0.037939 | 0.075879 | 0.25638 | 0.51276
19 | Ace Bypass | Exclude | 0.454 | 71.2309 | 10.1375 | 8.1942 | 7.0265 | -26.6436 | -5.2955 | ran before the background settled; | 0.037856 | 0.075712 | 0.25627 | 0.51253
20 | Random | Sample |  |  |  | 8.1732 | 7.0085 | -26.7065 | -5.1551 | d15N outside anchor range; ran before the background settled; | 0.037803 | 0.075605 | 0.25545 | 0.51089
21 | Random | Sample |  |  |  | 9.3775 | 8.0412 | -26.6175 | -5.2939 | d15N outside anchor range; ran before the background settled; | 0.037878 | 0.075757 | 0.25626 | 0.51251
22 | Ace Drift4 | Drift | 0.402 | 72.9394 | 10.3976 | 8.1808 | 7.015 | -26.6393 | -5.1812 | ran before the background settled; | 0.03786 | 0.075719 | 0.2556 | 0.5112
23 | Random | Sample |  |  |  | 0.45939 | 0.39393 | -35.1409 | -2.4733 | N above linearity range; C below linearity range; d15N outside anchor range; ran before the background settled; | 0.038489 | 0.076978 | 0.24043 | 0.48087
24 | Random | Sample |  |  |  | 8.2836 | 7.1032 | -26.6116 | -5.3234 | d15N outside anchor range; ran before the background settled; | 0.037883 | 0.075767 | 0.25643 | 0.51286
25 | Random | Sample |  |  |  | 1.9872 | 1.704 | -0.73486 | 41.0067 | d15N outside anchor range; d13C outside anchor range; ran before the background settled; | 0.093129 | 0.18626 | 0.2995 | 0.59899
26 | Ace Drift5 | Drift | 0.4 | 73.2534 | 10.4087 | 8.2072 | 7.0377 | -26.5412 | -5.3465 | ran before the background settled; | 0.037944 | 0.075889 | 0.25657 | 0.51313
27 | Random | Sample |  |  |  | 1.9833 | 1.7007 | -40.7659 | 1.7388 | d15N outside anchor range; | 0.046602 | 0.093204 | 0.21982 | 0.43965
28 | Random | Sample |  |  |  | 5.3432 | 4.5818 | -26.3982 | -5.4866 | d15N outside anchor range; | 0.038072 | 0.076143 | 0.25739 | 0.51478
29 | Random | Sample |  |  |  | 7.9356 | 6.8047 | -26.8321 | -4.8904 | C above linearity range; d15N outside anchor range; | 0.037698 | 0.075397 | 0.25391 | 0.50782
30 | Ace Drift6 | Drift | 0.418 | 72.4229 | 10.2194 | 8.2645 | 7.0868 | -26.6069 | -5.3845 |  | 0.037887 | 0.075775 | 0.25679 | 0.51358
31 | Random | Sample |  |  |  | 8.3225 | 7.1366 | -26.7535 | -5.0141 | d15N outside anchor range; | 0.037763 | 0.075526 | 0.25463 | 0.50925
32 | Random | Sample |  |  |  | 8.5323 | 7.3164 | -26.7013 | -5.5595 | d15N outside anchor range; | 0.037807 | 0.075614 | 0.25782 | 0.51564
33 | Random | Sample |  |  |  | 7.7611 | 6.6551 | -26.7622 | -4.9602 | N above linearity range; C above linearity range; d15N outside anchor range; | 0.037756 | 0.075512 | 0.25431 | 0.50862
34 | Ace Drift7 | Drift | 0.399 | 71.8523 | 10.4394 | 8.0266 | 6.8828 | -26.5959 | -5.3475 |  | 0.037897 | 0.075794 | 0.25657 | 0.51314
35 | Random | Sample |  |  |  | 8.0969 | 6.9431 | -26.7374 | -5.1476 | d15N outside anchor range; | 0.037777 | 0.075553 | 0.2554 | 0.5108
36 | Random | Sample |  |  |  | 8.1163 | 6.9597 | -26.6262 | -5.2176 | d15N outside anchor range; | 0.037871 | 0.075742 | 0.25581 | 0.51162
37 | Random | Sample |  |  |  | 7.9857 | 6.8477 | -26.72 | -4.9828 | d15N outside anchor range; | 0.037791 | 0.075582 | 0.25444 | 0.50889
38 | Ace Drift8 | Drift | 0.394 | 72.7514 | 10.5719 | 8.0252 | 6.8816 | -26.5628 | -5.5056 |  | 0.037926 | 0.075851 | 0.2575 | 0.515
39 | Random | Sample |  |  |  | 8.2577 | 7.081 | -26.5943 | -5.3026 | d15N outside anchor range; | 0.037898 | 0.075797 | 0.25631 | 0.51262
40 | Random | Sample |  |  |  | 41.3957 | 35.4968 | -27.7533 | 2.5055 | N below linearity range; C below linearity range; N peak only 1.5x the blank; | 0.037033 | 0.074066 | 0.21653 | 0.43306
41 | Random | Sample |  |  |  | 17.282 | 14.8193 | -27.801 | 1.4417 | N below linearity range; C below linearity range; d15N outside anchor range; N peak only 2.3x the blank; | 0.037003 | 0.074007 | 0.22114 | 0.44228
42 | Ace Drift9 | Drift | 0.38 | 71.679 | 10.4097 | 8.0301 | 6.8858 | -26.6476 | -5.254 |  | 0.037853 | 0.075705 | 0.25602 | 0.51205
43 | Random | Sample |  |  |  | 23.3962 | 20.0622 | -27.2996 | -0.47444 | N below linearity range; C below linearity range; d15N outside anchor range; N peak only 1.6x the blank; | 0.037339 | 0.074677 | 0.23015 | 0.4603
44 | Random | Sample |  |  |  | 62.0418 | 53.2008 | -27.7084 | 3.1358 | N below linearity range; C below linearity range; N peak only 1.3x the blank; | 0.037061 | 0.074123 | 0.21394 | 0.42789
45 | Random | Sample |  |  |  | 61.0199 | 52.3245 | -27.897 | 3.2249 | N below linearity range; C below linearity range; N peak only 1.3x the blank; | 0.036945 | 0.073891 | 0.21359 | 0.42717
46 | Ace Drift10 | Drift | 0.419 | 72.7665 | 10.6693 | 7.9536 | 6.8202 | -26.5787 | -5.1842 |  | 0.037912 | 0.075824 | 0.25562 | 0.51123
47 | Random | Sample |  |  |  | 61.9571 | 53.1282 | -27.6757 | 2.8877 | N below linearity range; C below linearity range; N peak only 1.3x the blank; | 0.037082 | 0.074164 | 0.21495 | 0.4299
48 | Random | Sample |  |  |  | 7.9987 | 6.8589 | -26.752 | -5.1734 | d15N outside anchor range; | 0.037764 | 0.075529 | 0.25555 | 0.51111
49 | Random | Sample |  |  |  | 7.8429 | 6.7253 | -26.7563 | -4.8718 | d15N outside anchor range; | 0.037761 | 0.075522 | 0.2538 | 0.5076
50 | Ace Drift11 | Drift | 0.424 | 70.6621 | 10.4051 | 7.9196 | 6.7911 | -26.6673 | -5.0518 |  | 0.037836 | 0.075671 | 0.25484 | 0.50969
51 | Random | Sample |  |  |  | 7.9494 | 6.8166 | -26.6941 | -4.9832 | d15N outside anchor range; | 0.037813 | 0.075626 | 0.25445 | 0.50889
52 | Random | Sample |  |  |  | 7.8494 | 6.7309 | -26.7448 | -4.7361 | d15N outside anchor range; | 0.03777 | 0.075541 | 0.25302 | 0.50603
53 | Random | Sample | [6 empty] | -8.2771 | no C peak; N below linearity range; d15N outside anchor range; N peak only 0.7x the blank; |  |  | 0.2744 | 0.54881
54 | Ace Drift12 | Drift | 0.38 | 72.5685 | 10.8905 | 7.7708 | 6.6635 | -26.5763 | -5.1308 |  | 0.037914 | 0.075828 | 0.2553 | 0.51061
55 | Ace Drift 13 | Drift | 0.397 | 72.4828 | 10.9555 | 7.7156 | 6.6161 | -26.6186 | -5.2162 |  | 0.037877 | 0.075755 | 0.2558 | 0.51161
56 | blank capsule | Blank | 0.397 |  | 0.030431 |  |  |  | -5.5933 | no C peak; N below linearity range; |  |  | 0.25802 | 0.51603

## QC
QC
Read this before taking any number off the Results sheet.
Charts sit beside the section they belong to. A trendline is Excel's own linear fit through the plotted points; for the drift monitors in the fit it is the row-model slope Fits reports.
Run
Autosampler positions detected | 56
IRMS peak records read from Raw | 386
Unknowns (role Sample) | 26
Positions with no nitrogen peak | 1
Working gas, from the reference pulses Isodat did not designate
The designated pulse carries the assigned value by definition, so it can never move. An export cut down to the sample peaks carries no pulses at all, and this section then says so rather than reporting a number.
N2 pulses read | 56 |  | ok
N2 pulse SD (permil) | 0.066341
N2 pulse trend, permil per position | -0.00084699 |  | ok, no trend (p=0.12)
N2 pulse range (permil) | 0.285
CO2 pulses read | 56 |  | ok
CO2 pulse SD (permil) | 0.065478
CO2 pulse trend, permil per position | 0.00088161 |  | ok, no trend (p=0.10)
CO2 pulse range (permil) | 0.422
Calibration
Nitrogen anchors used | 6 |  | ok
Nitrogen calibration slope | 1.0266 |  | ok
Carbon anchors used | 6 |  | ok
Carbon calibration slope | 1.0142 |  | ok
Check standards measured in this run | 0 |  | CHECK: nothing independent of the anchors tests this calibration; run a check standard (USGS66, or an in-house material with an assigned value) in every sequence
Samples outside the anchor d15N range | 22 |  | CHECK: these values are extrapolated beyond the anchors; add an anchor that brackets them
Samples outside the anchor d13C range | 1 |  | CHECK: these values are extrapolated beyond the anchors; add an anchor that brackets them
Is each correction supported by this run?
A slope is applied on Auto only when p is below the significance level in Setup B15. The FIRMS guide (section 6.4) asks that a correction be shown necessary before it is applied, because over-correction adds as much error as it removes.
Nitrogen linearity slope, permil per area unit | -0.056964 |  | slope -0.05696 +/- 0.01494, p = 0.062, not applied
Carbon linearity slope | 0.0017328 |  | slope 0.00173 +/- 0.00025, p = 0.020, applied
Nitrogen drift slope, per position or per background unit | -0.014754 |  | slope -0.01475 +/- 0.00271, p = 0.000, applied
Carbon drift slope | 0.0025977 |  | slope 0.00260 +/- 0.00085, p = 0.011, applied
With five or six standards the test has little power: a slope that fails it is unproven, not disproven. Set the switch to On to apply it regardless.
Nitrogen drift monitors flagged as outliers | 0 |  | ok
Carbon drift monitors flagged as outliers | 0 |  | ok
Nitrogen model in use | Row
Residual SD, row model (permil) | 0.14264
Residual SD, background model (permil) | 0.097601
Carbon model in use | Row
Residual SD, row model (permil) | 0.044946
Residual SD, background model (permil) | 0.043144
Residual SD is the scatter left in the drift standards themselves, so it cannot decide between the two models. The check standard offsets below can.
Blank
Blank correction in force, nitrogen | No |  | not applied: fewer than three blanks
Blank capsules with a nitrogen peak | 2
SD of blank d15N (permil) | 15.9509
Blank correction in force, carbon | No
Blank capsules with a carbon peak | 0
SD of blank d13C (permil)
A blank that gives no peak at all has not been measured, only bounded by the detection limit. The guide asks that its size and d value be determined, which the two-standard decreasing-size series does.
Check standards and drift monitors
Material | Role | n | mean d15N | accepted | offset | mean d13C | accepted | offset
USGS 64 | Anchor | 3 | 1.7621 | 1.76 | 0.0021172 | -40.81 | -40.81 | 2.8219e-05
USGS 65 | Anchor | 3 | 20.6779 | 20.68 | -0.0021172 | -20.29 | -20.29 | -2.8219e-05
USGS 66 | Check | 0 |  | 40.83 |  |  | -0.67
USGS 40 | Check | 0 |  | -4.52 |  |  | -26.39
USGS 41a | Check | 0 |  | 47.55 |  |  | 36.55
Ace Drift | Drift | 13 | -5.2376 |  |  | -26.6053
An anchor is forced onto its accepted value by construction, so only the check rows are an independent test. Offsets are in per mil.
Uncertainty
Sample repeatability in use, d15N (permil) | 0.16882 |  | pooled anchor replicates
Sample repeatability in use, d13C (permil) | 0.020549 |  | pooled anchor replicates
Coverage factor k | 2
Typical u(d15N) of the unknowns (permil) | 0.24729
Typical u(d13C) of the unknowns (permil) | 0.040232
u combines the certificate uncertainty of each anchor, the scatter of its replicates, and the sample repeatability, through the fitted line, as a Kragten spreadsheet would. It leaves out the blank, linearity and drift corrections and the day-to-day reproducibility a control chart would give, so it is a floor, not the whole budget.
SD of the drift standards, d15N (permil) | 0.1402
SD of the drift standards, d13C (permil) | 0.043641
These understate the real precision, because the drift standards are the material the drift correction was fitted to. The check standard offsets above are the honest test.
Flags
Rows carrying a flag on Results | 48
Minimum trusted Area 28 | 8.889
Maximum trusted Area 28 | 51.069
Minimum trusted Area 44 | 37.895
Maximum trusted Area 44 | 240.154
Samples outside the linearity range | 14 |  | CHECK: weigh samples to land inside the standards, or extend the size series
Instrument settling
A run that began before the mass-28 background reached its floor carries a settling transient, not a steady drift, and a straight line in position is the wrong shape for it.
Reference-peak background floor | 23.1
Starting excess above the floor | 11.1
Background settled from position | 27 |  | CHECK: positions 1 to 26 ran before the background settled
First position used in the drift fit | all |  | consider setting Setup B14 to 27
Linearity standards before that position | 4 |  | CHECK: size and time are confounded, the linearity fit sits inside the unsettled window
Calibration anchors before that position | 6 |  | CHECK: the scale was set while the instrument was still settling
Drift slope before the settled position | -0.0318
Drift slope from the settled position | -0.0028939 |  | settling transient, not steady drift
Sample-peak background on the drift monitors, lowest to highest | 23.8 to 30.3
Its trend, per position | -0.11963 |  | CHECK: the background under the monitors moved across the run (p=0.000)
The drift monitors are all one size, so their background follows the instrument.
A high background right after a large sample points to carryover from that sample, not to instrument drift.
Does each correction reduce the scatter of the drift standards?
as measured | 0.26271
after blank | 0.26271
after linearity | 0.26271
after drift | 0.13657
after calibration | 0.1402
SD in per mil.
This table can judge the blank and drift corrections but not the linearity one, because the drift standards are all about the same size and a size-dependent correction moves them together. Judge linearity by the check standards and by the spread of the linearity series itself.
Carryover after oversized samples
A sample above the linearity range can leave something in the EA that raises the background of the next one or two analyses.
Positions above the linearity range | 4
Of those, followed directly by a sample or check | 1 |  | CHECK: put a blank capsule or a drift monitor after each oversized sample
Positions flagged for possible carryover | 0 |  | ok
Flagged position | Position | Rise (mV) | Share of peak height
The rise is measured against the last background before the oversized samples, so a slow change across the run does not count. The same rise weighs more under a small peak.